METRUM INFORMATION STORAGE LTD. - Wells, England. BA5 1BB

PROJECT NUMBER	1098
DOCUMENT NUMBER	1686030020
ISSUE	Issue 1
SCN	N/A
DATE	12/02/98

FACTORY ACCEPTANCE TEST PROCEDURE 
FOR THE
EVN/Mark IV STATION UNIT PHASE CAL

APPROVAL:

AUTHOR. .......................................................................P Hazell/ K.Hosking

METRUM:.......................................................................M Clemow

JIVE:...............................................................................B Anderson

This specification contains 58 pages.

This specification (including the copyright therein) is owned by METRUM INFORMATION STORAGE Ltd. All information contained therein must be kept confidential and used solely for the purposes specified by the owner. This specification must not be copied or reproduced nor divulged in whole or part without the owners prior written consent.

# ii	TEST RECORD

BOARD SERIAL NUMBER

DATE TESTING COMMENCED

DATE TESTING COMPLETED

PLACE OF TEST

REMARKS:

| FOR CUSTOMER | ........................................... | (SIGNATURE) |
| --- | --- | --- |
|  | ........................................... | (NAME) |
| FOR METRUM INFORMATION STORAGE LTD. | ........................................... | (SIGNATURE) |
|  | ........................................... | (NAME) |

#### iii	TEST RESULTS SUMMARY

| Test Number | Details of Failure | Corrective Action | Re-test |  |
| --- | --- | --- | --- | --- |
|  |  |  | Date | Pass/Fail |

ii TEST RECORD

1 INTRODUCTION

2 EQUIPMENT AND FACILITIES

2.1 Facilities

2.2 Power Supplies

2.3 Test Equipment

2.3.1 Calibration

2.3.2 Test Equipment

2.3.3 Test Set Up

2.4 Test Software

2.4.1 SU Software

2.4.2 C3 Emulation

2.4.3 PCM Simulation

3 ACCEPTANCE TESTS.

3.1 General Instructions

3.2 Safety and Quality Checks

3.3 Firmware

3.4 SUCC Initialisation

4 PCM TESTS.

M32 1_4

32meg 1_4 00

32meg 1_4 53

32meg 1_4 5E

32meg 1_4 5F

32meg 1_4 6F

32meg 1_4 7F

32meg 1_4 A6

32meg 1_4 AD

32meg 1_4 AF

32meg 1_4 B7

32meg 1_4 BF

32meg 1_4 DB

32meg 1_4 DE

32meg 1_4 DF

32meg 1_4 EF

32meg 1_4 FF

M16 1_2

16meg 1_2 FF

16meg 1_2 00

16meg 1_2 11

16meg 1_2 22

M16 M08 M04 1_1

16meg 08meg 04meg 1_1 FF

M08 M04 M02 2_1

02meg 2_1 FF / 04meg 2_1 FF / 08meg 2_1 FF

02meg 2_1 00 / 04meg 2_1 00 / 08meg 2_1 00

M04 M02 4_1

02meg 4_1 FF / 04meg 4_1 FF

Random Table data and Stress tests for large counts.

32meg 1_4 00 Random Data

32meg 1_4 53 Random Data

32meg 1_4 5E Random Data

32meg 1_4 5F Random Data

32meg 1_4 6F Random Data

32meg 1_4 7F Random Data

32meg 1_4 A6 Random Data

32meg 1_4 AD Random Data

32meg 1_4 AF Random Data

32meg 1_4 B7 Random Data

32meg 1_4 BF Random Data

32meg 1_4 DB Random Data

32meg 1_4 DE Random Data

32meg 1_4 DF Random Data

32meg 1_4 EF Random Data

32meg 1_4 FF Random Data

Tests for non zero data in the validity equals zero table section.

32meg 1_4 00 Non Zero V=0 Table Data

32meg 1_4 53 Non Zero V=0 Table Data

32meg 1_4 5E Non Zero V=0 Table Data

2meg 1_4 5F Non Zero V=0 Table Data

32meg 1_4 6F Non Zero V=0 Table Data

32meg 1_4 7F Non Zero V=0 Table Data

32meg 1_4 A6 Non Zero V=0 Table Data

32meg 1_4 AD Non Zero V=0 Table Data

32meg 1_4 AF Non Zero V=0 Table Data

32meg 1_4 B7 Non Zero V=0 Table Data

32meg 1_4 BF Non Zero V=0 Table Data

32meg 1_4 DB Non Zero V=0 Table Data

32meg 1_4 DE Non Zero V=0 Table Data

32meg 1_4 DF Non Zero V=0 Table Data

32meg 1_4 EF Non Zero V=0 Table Data

32meg 1_4 FF Non Zero V=0 Table Data

5 Notes For Test Method.

6 PCAL Table Contents.

7 CRM Output Patterns For Testing The PCM.

8 PCMSIM Program

#### i	AMENDMENT RECORD

| Date | Issue | Description |
| --- | --- | --- |
| 12th February 1998 | 1 | First Issue |
| 1st June 1999 | 1 | Details of Firmware changed on page 12 (SCN 10024) |

# 1	INTRODUCTION

There are two objectives to Phase Cal testing.  First to demonstrate that the Xilinx accumulators function in all required respects.  Second to show that the SUCC and Phase Cal operational software performs as intended.

To confirm correct accumulator operation we use essentially the same test software as we have in the development in both the SUCC and the PCM boards.  However it is resident with the operational code.  The advantage of this is that exactly the same XILINX configuration is used in both cases.  The functional selection is performed at power up on the PCM board which in turn signals the SUCC to function in either test or operational mode.  The test configuration gives a simple direct control over the number of frames over which to accumulate.  The operational code has features which reduce this direct control.

Once correct operation of the XILINX is shown a few representative operational tests will be performed.

For an introduction to the testing strategy please refer to the sections “Notes for test method” and “CRM output patterns for testing the PCM”.

The following directory structure is proposed to contain the PCM fat tests.

Directory PATH listing for Volume MS-DOS_5   Volume Serial Number is 1F88-4938 C:PCALTEST

+---TEMP

+---M32

¦	+---1_4

¦	¦	+---CRMDAT32

¦	+---1_2

¦		+---CRMDAT32

+---M16

¦	+---1_4

¦	¦	+---CRMDAT16

¦	+---1_2

¦	¦	+---CRMDAT16

¦	+---1_1

¦		+---CRMDAT16

+---M08

¦	+---1_2

¦	¦	+---CRMDAT08

¦	+---1_1

¦	¦	+---CRMDAT08

¦	+---2_1

¦		+---CRMDAT08

+---M04

¦	+---1_1

¦	¦	+---CRMDAT04

¦	+---2_1

¦	¦	+---CRMDAT04

¦	+---4_1

¦		+---CRMDAT04

+---M02

+---2_1

¦	+---CRMDAT02

+---4_1

The directory structure is

c:\pcaltest\ Mxx \mux\data.

Where Mxx gives the nominal frequency at which the Phase Cal is running. xx has values 32,16,08,04,02.

Under each frequency directory are all the possible mux ratios.

1_4	That is create one channel from 4 tracks.

1_2	That is create one channel from 2 tracks.

1_1	That is create one channel from 4 tracks.

2_1	That is create two channels from 1 track.

4_1	That is create four channels from 1 track.

A set of batch files are used to run the tests in a similar manner to the full SU FAT.

Under the mux directory is a data directory containing the CRM configurations

The 1_4 case provides the greatest variety of wave forms as shown in the appendix.  The 1_2 case can provide a 0101 or 1010 pattern.  The others provide either all zeros or all ones.

# 2	EQUIPMENT AND FACILITIES

## 2.1	Facilities

## 2.2	Power Supplies

240 Vac 50 Hz single phase outlets are required to power the system under test and test equipment. All supplies are to be from the same phase.

## 2.3	Test Equipment

### 2.3.1	Calibration

All measurement equipment used in the performance of these tests must either carry a current calibration label traceable to national standards or in the case of variable range equipment's, be calibrated locally using instruments which have been calibrated to national standards.

### 2.3.2	Test Equipment

PC with the following installed and configured (PCCC):

- NE2000 compatible ethernet card, BNC output.

- PC/TCPR Network software kernel for DOS 4.0

- AUTOEXEC.BAT and CONFIG.SYS files set up for tftp server.

- RAM.HEX  - file to be downloaded to SUCC

- TFTPSERV.BAT - batch file to set-up PC as TFTP server.

2 off BNC “T” pieces

2 off BNC 50 Ohm termination

1 off BNC - AUI transceiver e.g. Allied Telesis model AT-MX10

BNC - BNC, RG58 coax cable

Dumb Terminal

### 2.3.3	Test Set Up

Fit the AUI transceiver to the Ethernet connector on the MVME712M Transition Module.

Using the coaxial cable connect the PC BNC Ethernet to the AUI module via the “T” pieces with 50 Ohm terminators at each end.

Power up the PC, select TFTP server configuration from the boot menu (this will be the default which will come up after a few seconds).

## 2.4	Test Software

### 2.4.1	SU Software

The UUT is loaded with its normal operational software during testing.

(Host 192.9.200.7 c:\fatp\ram.hex  Issue:succ4114)

### 2.4.2	C3 Emulation

The C3 is emulated for testing on a PC (known as the PCCC) using SUTEST.EXE

This programme issues commands and receives responses via the Ethernet, TCP/IP interface in accordance with EVN42.

The format of tests is:

SUTEST <menu_file_name>

Menu file actions are simple instructions to cause SUTEST.EXE to perform the following operations.

OPEN		Open a socket to a SU.

CLOSE	Close the currently open socket.

SEND		Transmit a data file to the SU.

RECV		Receive data from the SU into a file.

LOOP		Loop waiting for user to press a key.

The data files used for sending to the SU are ASCII HEX files containing the actual C3 command plus a small header to define the input data format. The received data files are built using simple formatting which produces byte, word or long word ASCII HEX per line.

In general, tests are further simplified by the use of DOS batch files which, directly or via other batch files, call SUTEST.EXE with a specified menu file.

### 2.4.3	PCM Simulation

The expected results tables have been calculated using the PCMSIM program.  See section PCMSIM PROGRAM for the source code.

# 3	ACCEPTANCE TESTS.

## 3.1	General Instructions

·	Provision is made to record test results within this document. Take a copy for this purpose and begin by entering the details requested on page 3. Record details of any failed test on page 4.

Unless otherwise directed indicate a passed test thus:

| PASS | FAIL |
| --- | --- |

## 3.2	Safety and Quality Checks

Not applicable.

## 3.3	Firmware

Confirm that for the boards used the correct programmable devices are fitted:-

| Module and No. fitted | DEVICE | Part NO. | Issue |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| SUCC (x1) | XU1, XU2 | 1690020500 | A |  |  |  |
| DIM (x2) | U21 | 1697020000 | A |  |  |  |
| TRM (x8) | U33, U34 | 1690040005 | C (TRM4143) |  |  |  |
|  | U38 | 1695040005 | A |  |  |  |
|  | U178* | 1696040005 | A |  |  |  |
|  | U40 | 1697040005 | A |  |  |  |
|  | U39 | 1698040005 | A |  |  |  |
| CRM (x1) | U42, U43 | 1690020010 | D (CRM4135) |  |  |  |
|  | U48 | 1695020010 | A |  |  |  |
|  | U37 | 1697020010 | A |  |  |  |
|  | U52 | 1698020010 | A |  |  |  |
| PCM (x1) | U54, U55 | 1690040020 | A (PCM4131) |  |  |  |
|  | U57 | 1697020020 | A |  |  |  |
| DCM (x1) | U19, U20 | 1690020015 | F (DCM4137) |  |  |  |
|  | U51 | 1695020015 | A |  |  |  |
|  | U38 | 1697020015 | A |  |  |  |
| DMM (x2) | U23 | 1695020016 | A |  |  |  |
| SUIM (x1) | N/A | N/A | N/A |  |  |  |
|  | *production TRM only |  |  |  |  |  |
|  |  |  |  | PASS |  | FAIL |

## 3.4	SUCC Initialisation

Connect a dumb terminal to the Console port of the MVME 712  (set-up: 9600/8/n/1)

Apply power to the Unit Under Test

Hit a key on the terminal keyboard within 2 seconds.

Type M for modify and select the following:

Mode 4 [run ftp boot loader]

Do you want a LAN I/F [Y]

This board’s IP address [...] set to 192.9.200.7

Subnet mask [N]

Shared memory network interface [N]

Default gateway [N]

Multiprocessor system [N]

Baud rate for serial channels [9600]

Bus address [1000000]

IP address of TFTP boot loader [192.9.200.130]

File name... [c:\fatp\ram.hex ]

How long to delay... [2]

Then chose “C” to continue.

Observe the Server screen and confirm that the following messages are displayed:

a) PC/TCP TFTP server initialised

b) Host 192.9.200.7 reading file c:\fatp\ram.hex  octet

.............

transfer successful

? bytes in ? seconds ?baud

Observe the DIM LED’s and confirm that they sequence through red, yellow, green and off.

|  |  | PASS | FAIL |
| --- | --- | --- | --- |

# 4	PCM TESTS.

We have set out a number of tests with sparse SRAM tables.  The table content is detailed the section PCAL Table Contents below. We use operational mode software to accumulate for one more frame than the number of frames specified in the Phase Cal request from the C3.  This departs from EVN 42 in that the message specification defines the nominal number of seconds to accumulate over.  Further we set the most significant bit in the long word to force collection over a single period.  In general we specify 3(+1) frames.  So that the total number of samples in the 1_4 case is 2500*8*4*4 = 320000 samples.

For the first set of tests all Table rams are filled with the same data.  This should result in all even channels accumulating the same data as channel 0.  These are out lined in the count tables for count0 through 8.  The due to a different multiplex arrangement in the 1_4 case for odd channels in the CRM the counts are not the same as for the even channels.  It is possible to determine the expected counts for odd channels by using the predicted results for the relevant wave form.

## M32 1_4

For 32meg 1_4 and a CRM table pattern of FF.  We expect to get S=M=V=1 for all data samples, all of which are valid, over a 4 frame test we expect  2500*8*4*4 = 3200000[4E200] for validity. The modulo 3200 counter loops 3200000/3200 = 100 times. With SRAM data from TABT000.DAT which has a single bit set per counter we expect 100[64] for each count.  These counts in themselves do not prove the correct position relationship due to the cyclic nature of the test.  We do however (at least during development) show on  the logic analyser trace the position of the low order data bit in the expected place.  Using SRAM tables TABT001.DAT and TABT002.DAT  for 32meg 1_4 and a CRM table pattern of FF we get the same counts.

With SRAM data from TABT003, the SRAM tables are the inverse of TABT000.DAT, so over a 4 frame test we expect 3200000[4E200] for validity and 319900[4E19C] for each count.  The following tables give the expected results for TABT000.  The TABT003 results may be inferred...

With the 32meg 1_4 condition we can apply the following a CRM table patterns:-

Run “C:>PCAL32  \PCALDATA\TABT000.DAT 32”     for TABT000.

Check that the results agree with the TABT000.DAT column of the following tables.

Then run “C:>PCAL32  \PCALDATA\TABT003.DAT 32”     for TABT003.

Check that the results agree with the TABT003.DAT column of the following tables.

Results are stored in 32**.RSP where ** refers to the CRM configuration e.g. FF, 53 etc.  The results files are overwritten each time the test is run.

### 32meg 1_4 00

00000000 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000000 | 0x00000000 |  |  |  |
| count1 | 0x00000000 | 0x00000000 |  |  |  |
| count2 | 0x00000000 | 0x00000000 |  |  |  |
| count3 | 0x00000000 | 0x00000000 |  |  |  |
| count4 | 0x00000000 | 0x00000000 |  |  |  |
| count5 | 0x00000000 | 0x00000000 |  |  |  |
| count6 | 0x00000000 | 0x00000000 |  |  |  |
| count7 | 0x00000000 | 0x00000000 |  |  |  |
| count8 | 0x00000000 | 0x00000000 |  |  |  |
| count0 | 0x00000000 | 0x00000000 |  |  |  |
| count1 | 0x00000000 | 0x00000000 |  |  |  |
| count2 | 0x00000000 | 0x00000000 |  |  |  |
| count3 | 0x00000000 | 0x00000000 |  |  |  |
| count4 | 0x00000000 | 0x00000000 |  |  |  |
| count5 | 0x00000000 | 0x00000000 |  |  |  |
| count6 | 0x00000000 | 0x00000000 |  |  |  |
| count7 | 0x00000000 | 0x00000000 |  |  |  |
| count8 | 0x00000000 | 0x00000000 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 32meg 1_4 53

00010001 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000000 | 0x00013880 |  |  |  |
| count1 | 0x00000000 | 0x00013880 |  |  |  |
| count2 | 0x00000000 | 0x00013880 |  |  |  |
| count3 | 0x00000064 | 0x0001381C |  |  |  |
| count4 | 0x00000000 | 0x00013880 |  |  |  |
| count5 | 0x00000000 | 0x00013880 |  |  |  |
| count6 | 0x00000000 | 0x00013880 |  |  |  |
| count7 | 0x00000064 | 0x0001381C |  |  |  |
| count8 | 0x00013880 | 0x00013880 |  |  |  |
| count0 | 0x00000000 | 0x00013880 |  |  |  |
| count1 | 0x00000064 | 0x0001381C |  |  |  |
| count2 | 0x00000000 | 0x00013880 |  |  |  |
| count3 | 0x00000000 | 0x00013880 |  |  |  |
| count4 | 0x00000000 | 0x00013880 |  |  |  |
| count5 | 0x00000064 | 0x0001381C |  |  |  |
| count6 | 0x00000000 | 0x00013880 |  |  |  |
| count7 | 0x00000000 | 0x00013880 |  |  |  |
| count8 | 0x00013880 | 0x00013880 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 32meg 1_4 5E

01000100 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000000 | 0x00013880 |  |  |  |
| count1 | 0x00000064 | 0x0001381C |  |  |  |
| count2 | 0x00000000 | 0x00013880 |  |  |  |
| count3 | 0x00000000 | 0x00013880 |  |  |  |
| count4 | 0x00000000 | 0x00013880 |  |  |  |
| count5 | 0x00000064 | 0x0001381C |  |  |  |
| count6 | 0x00000000 | 0x00013880 |  |  |  |
| count7 | 0x00000000 | 0x00013880 |  |  |  |
| count8 | 0x00013880 | 0x00013880 |  |  |  |
| count0 | 0x00000000 | 0x00013880 |  |  |  |
| count1 | 0x00000000 | 0x00013880 |  |  |  |
| count2 | 0x00000000 | 0x00013880 |  |  |  |
| count3 | 0x00000064 | 0x0001381C |  |  |  |
| count4 | 0x00000000 | 0x00013880 |  |  |  |
| count5 | 0x00000000 | 0x00013880 |  |  |  |
| count6 | 0x00000000 | 0x00013880 |  |  |  |
| count7 | 0x00000064 | 0x0001381C |  |  |  |
| count8 | 0x00013880 | 0x00013880 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 32meg 1_4 5F

01010101 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000000 | 0x00027100 |  |  |  |
| count1 | 0x00000064 | 0x0002709C |  |  |  |
| count2 | 0x00000000 | 0x00027100 |  |  |  |
| count3 | 0x00000064 | 0x0002709C |  |  |  |
| count4 | 0x00000000 | 0x00027100 |  |  |  |
| count5 | 0x00000064 | 0x0002709C |  |  |  |
| count6 | 0x00000000 | 0x00027100 |  |  |  |
| count7 | 0x00000064 | 0x0002709C |  |  |  |
| count8 | 0x00027100 | 0x00027100 |  |  |  |
| count0 | 0x00000000 | 0x00027100 |  |  |  |
| count1 | 0x00000064 | 0x0002709C |  |  |  |
| count2 | 0x00000000 | 0x00027100 |  |  |  |
| count3 | 0x00000064 | 0x0002709C |  |  |  |
| count4 | 0x00000000 | 0x00027100 |  |  |  |
| count5 | 0x00000064 | 0x0002709C |  |  |  |
| count6 | 0x00000000 | 0x00027100 |  |  |  |
| count7 | 0x00000064 | 0x0002709C |  |  |  |
| count8 | 0x00027100 | 0x00027100 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 32meg 1_4 6F

01100110 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000000 | 0x00027100 |  |  |  |
| count1 | 0x00000064 | 0x0002709C |  |  |  |
| count2 | 0x00000064 | 0x0002709C |  |  |  |
| count3 | 0x00000000 | 0x00027100 |  |  |  |
| count4 | 0x00000000 | 0x00027100 |  |  |  |
| count5 | 0x00000064 | 0x0002709C |  |  |  |
| count6 | 0x00000064 | 0x0002709C |  |  |  |
| count7 | 0x00000000 | 0x00027100 |  |  |  |
| count8 | 0x00027100 | 0x00027100 |  |  |  |
| count0 | 0x00000064 | 0x0002709C |  |  |  |
| count1 | 0x00000000 | 0x00027100 |  |  |  |
| count2 | 0x00000000 | 0x00027100 |  |  |  |
| count3 | 0x00000064 | 0x0002709C |  |  |  |
| count4 | 0x00000064 | 0x0002709C |  |  |  |
| count5 | 0x00000000 | 0x00027100 |  |  |  |
| count6 | 0x00000000 | 0x00027100 |  |  |  |
| count7 | 0x00000064 | 0x0002709C |  |  |  |
| count8 | 0x00027100 | 0x00027100 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 32meg 1_4 7F

01110111 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000000 | 0x0003A980 |  |  |  |
| count1 | 0x00000064 | 0x0003A91C |  |  |  |
| count2 | 0x00000064 | 0x0003A91C |  |  |  |
| count3 | 0x00000064 | 0x0003A91C |  |  |  |
| count4 | 0x00000000 | 0x0003A980 |  |  |  |
| count5 | 0x00000064 | 0x0003A91C |  |  |  |
| count6 | 0x00000064 | 0x0003A91C |  |  |  |
| count7 | 0x00000064 | 0x0003A91C |  |  |  |
| count8 | 0x0003A980 | 0x0003A980 |  |  |  |
| count0 | 0x00000064 | 0x0003A91C |  |  |  |
| count1 | 0x00000064 | 0x0003A91C |  |  |  |
| count2 | 0x00000000 | 0x0003A980 |  |  |  |
| count3 | 0x00000064 | 0x0003A91C |  |  |  |
| count4 | 0x00000064 | 0x0003A91C |  |  |  |
| count5 | 0x00000064 | 0x0003A91C |  |  |  |
| count6 | 0x00000000 | 0x0003A980 |  |  |  |
| count7 | 0x00000064 | 0x0003A91C |  |  |  |
| count8 | 0x0003A980 | 0x0003A980 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 32meg 1_4 A6

00100010 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000000 | 0x00013880 |  |  |  |
| count1 | 0x00000000 | 0x00013880 |  |  |  |
| count2 | 0x00000064 | 0x0001381C |  |  |  |
| count3 | 0x00000000 | 0x00013880 |  |  |  |
| count4 | 0x00000000 | 0x00013880 |  |  |  |
| count5 | 0x00000000 | 0x00013880 |  |  |  |
| count6 | 0x00000064 | 0x0001381C |  |  |  |
| count7 | 0x00000000 | 0x00013880 |  |  |  |
| count8 | 0x00013880 | 0x00013880 |  |  |  |
| count0 | 0x00000064 | 0x0001381C |  |  |  |
| count1 | 0x00000000 | 0x00013880 |  |  |  |
| count2 | 0x00000000 | 0x00013880 |  |  |  |
| count3 | 0x00000000 | 0x00013880 |  |  |  |
| count4 | 0x00000064 | 0x0001381C |  |  |  |
| count5 | 0x00000000 | 0x00013880 |  |  |  |
| count6 | 0x00000000 | 0x00013880 |  |  |  |
| count7 | 0x00000000 | 0x00013880 |  |  |  |
| count8 | 0x00013880 | 0x00013880 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 32meg 1_4 AD

10001000 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000064 | 0x0001381C |  |  |  |
| count1 | 0x00000000 | 0x00013880 |  |  |  |
| count2 | 0x00000000 | 0x00013880 |  |  |  |
| count3 | 0x00000000 | 0x00013880 |  |  |  |
| count4 | 0x00000064 | 0x0001381C |  |  |  |
| count5 | 0x00000000 | 0x00013880 |  |  |  |
| count6 | 0x00000000 | 0x00013880 |  |  |  |
| count7 | 0x00000000 | 0x00013880 |  |  |  |
| count8 | 0x00013880 | 0x00013880 |  |  |  |
| count0 | 0x00000000 | 0x00013880 |  |  |  |
| count1 | 0x00000000 | 0x00013880 |  |  |  |
| count2 | 0x00000064 | 0x0001381C |  |  |  |
| count3 | 0x00000000 | 0x00013880 |  |  |  |
| count4 | 0x00000000 | 0x00013880 |  |  |  |
| count5 | 0x00000000 | 0x00013880 |  |  |  |
| count6 | 0x00000064 | 0x0001381C |  |  |  |
| count7 | 0x00000000 | 0x00013880 |  |  |  |
| count8 | 0x00013880 | 0x00013880 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 32meg 1_4 AF

10101010 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000064 | 0x0002709C |  |  |  |
| count1 | 0x00000000 | 0x00027100 |  |  |  |
| count2 | 0x00000064 | 0x0002709C |  |  |  |
| count3 | 0x00000000 | 0x00027100 |  |  |  |
| count4 | 0x00000064 | 0x0002709C |  |  |  |
| count5 | 0x00000000 | 0x00027100 |  |  |  |
| count6 | 0x00000064 | 0x0002709C |  |  |  |
| count7 | 0x00000000 | 0x00027100 |  |  |  |
| count8 | 0x00027100 | 0x00027100 |  |  |  |
| count0 | 0x00000064 | 0x0002709C |  |  |  |
| count1 | 0x00000000 | 0x00027100 |  |  |  |
| count2 | 0x00000064 | 0x0002709C |  |  |  |
| count3 | 0x00000000 | 0x00027100 |  |  |  |
| count4 | 0x00000064 | 0x0002709C |  |  |  |
| count5 | 0x00000000 | 0x00027100 |  |  |  |
| count6 | 0x00000064 | 0x0002709C |  |  |  |
| count7 | 0x00000000 | 0x00027100 |  |  |  |
| count8 | 0x00027100 | 0x00027100 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 32meg 1_4 B7

00110011 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000000 | 0x00027100 |  |  |  |
| count1 | 0x00000000 | 0x00027100 |  |  |  |
| count2 | 0x00000064 | 0x0002709C |  |  |  |
| count3 | 0x00000064 | 0x0002709C |  |  |  |
| count4 | 0x00000000 | 0x00027100 |  |  |  |
| count5 | 0x00000000 | 0x00027100 |  |  |  |
| count6 | 0x00000064 | 0x0002709C |  |  |  |
| count7 | 0x00000064 | 0x0002709C |  |  |  |
| count8 | 0x00027100 | 0x00027100 |  |  |  |
| count0 | 0x00000064 | 0x0002709C |  |  |  |
| count1 | 0x00000064 | 0x0002709C |  |  |  |
| count2 | 0x00000000 | 0x00027100 |  |  |  |
| count3 | 0x00000000 | 0x00027100 |  |  |  |
| count4 | 0x00000064 | 0x0002709C |  |  |  |
| count5 | 0x00000064 | 0x0002709C |  |  |  |
| count6 | 0x00000000 | 0x00027100 |  |  |  |
| count7 | 0x00000000 | 0x00027100 |  |  |  |
| count8 | 0x00027100 | 0x00027100 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 32meg 1_4 BF

10111011 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000064 | 0x0003A91C |  |  |  |
| count1 | 0x00000000 | 0x0003A980 |  |  |  |
| count2 | 0x00000064 | 0x0003A91C |  |  |  |
| count3 | 0x00000064 | 0x0003A91C |  |  |  |
| count4 | 0x00000064 | 0x0003A91C |  |  |  |
| count5 | 0x00000000 | 0x0003A980 |  |  |  |
| count6 | 0x00000064 | 0x0003A91C |  |  |  |
| count7 | 0x00000064 | 0x0003A91C |  |  |  |
| count8 | 0x0003A980 | 0x0003A980 |  |  |  |
| count0 | 0x00000064 | 0x0003A91C |  |  |  |
| count1 | 0x00000064 | 0x0003A91C |  |  |  |
| count2 | 0x00000064 | 0x0003A91C |  |  |  |
| count3 | 0x00000000 | 0x0003A980 |  |  |  |
| count4 | 0x00000064 | 0x0003A91C |  |  |  |
| count5 | 0x00000064 | 0x0003A91C |  |  |  |
| count6 | 0x00000064 | 0x0003A91C |  |  |  |
| count7 | 0x00000000 | 0x0003A980 |  |  |  |
| count8 | 0x0003A980 | 0x0003A980 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 32meg 1_4 DB

10011001 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000064 | 0x0002709C |  |  |  |
| count1 | 0x00000000 | 0x00027100 |  |  |  |
| count2 | 0x00000000 | 0x00027100 |  |  |  |
| count3 | 0x00000064 | 0x0002709C |  |  |  |
| count4 | 0x00000064 | 0x0002709C |  |  |  |
| count5 | 0x00000000 | 0x00027100 |  |  |  |
| count6 | 0x00000000 | 0x00027100 |  |  |  |
| count7 | 0x00000064 | 0x0002709C |  |  |  |
| count8 | 0x00027100 | 0x00027100 |  |  |  |
| count0 | 0x00000000 | 0x00027100 |  |  |  |
| count1 | 0x00000064 | 0x0002709C |  |  |  |
| count2 | 0x00000064 | 0x0002709C |  |  |  |
| count3 | 0x00000000 | 0x00027100 |  |  |  |
| count4 | 0x00000000 | 0x00027100 |  |  |  |
| count5 | 0x00000064 | 0x0002709C |  |  |  |
| count6 | 0x00000064 | 0x0002709C |  |  |  |
| count7 | 0x00000000 | 0x00027100 |  |  |  |
| count8 | 0x00027100 | 0x00027100 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 32meg 1_4 DE

11001100 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000064 | 0x0002709C |  |  |  |
| count1 | 0x00000064 | 0x0002709C |  |  |  |
| count2 | 0x00000000 | 0x00027100 |  |  |  |
| count3 | 0x00000000 | 0x00027100 |  |  |  |
| count4 | 0x00000064 | 0x0002709C |  |  |  |
| count5 | 0x00000064 | 0x0002709C |  |  |  |
| count6 | 0x00000000 | 0x00027100 |  |  |  |
| count7 | 0x00000000 | 0x00027100 |  |  |  |
| count8 | 0x00027100 | 0x00027100 |  |  |  |
| count0 | 0x00000000 | 0x00027100 |  |  |  |
| count1 | 0x00000000 | 0x00027100 |  |  |  |
| count2 | 0x00000064 | 0x0002709C |  |  |  |
| count3 | 0x00000064 | 0x0002709C |  |  |  |
| count4 | 0x00000000 | 0x00027100 |  |  |  |
| count5 | 0x00000000 | 0x00027100 |  |  |  |
| count6 | 0x00000064 | 0x0002709C |  |  |  |
| count7 | 0x00000064 | 0x0002709C |  |  |  |
| count8 | 0x00027100 | 0x00027100 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 32meg 1_4 DF

11011101 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000064 | 0x0003A91C |  |  |  |
| count1 | 0x00000064 | 0x0003A91C |  |  |  |
| count2 | 0x00000000 | 0x0003A980 |  |  |  |
| count3 | 0x00000064 | 0x0003A91C |  |  |  |
| count4 | 0x00000064 | 0x0003A91C |  |  |  |
| count5 | 0x00000064 | 0x0003A91C |  |  |  |
| count6 | 0x00000000 | 0x0003A980 |  |  |  |
| count7 | 0x00000064 | 0x0003A91C |  |  |  |
| count8 | 0x0003A980 | 0x0003A980 |  |  |  |
| count0 | 0x00000000 | 0x0003A980 |  |  |  |
| count1 | 0x00000064 | 0x0003A91C |  |  |  |
| count2 | 0x00000064 | 0x0003A91C |  |  |  |
| count3 | 0x00000064 | 0x0003A91C |  |  |  |
| count4 | 0x00000000 | 0x0003A980 |  |  |  |
| count5 | 0x00000064 | 0x0003A91C |  |  |  |
| count6 | 0x00000064 | 0x0003A91C |  |  |  |
| count7 | 0x00000064 | 0x0003A91C |  |  |  |
| count8 | 0x0003A980 | 0x0003A980 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 32meg 1_4 EF

11101110 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |
| --- | --- | --- |
| count0 | 0x00000064 | 0x0003A91C |
| count1 | 0x00000064 | 0x0003A91C |
| count2 | 0x00000064 | 0x0003A91C |
| count3 | 0x00000000 | 0x0003A980 |
| count4 | 0x00000064 | 0x0003A91C |
| count5 | 0x00000064 | 0x0003A91C |
| count6 | 0x00000064 | 0x0003A91C |
| count7 | 0x00000000 | 0x0003A980 |
| count8 | 0x0003A980 | 0x0003A980 |
| count0 | 0x00000064 | 0x0003A91C |
| count1 | 0x00000000 | 0x0003A980 |
| count2 | 0x00000064 | 0x0003A91C |
| count3 | 0x00000064 | 0x0003A91C |
| count4 | 0x00000064 | 0x0003A91C |
| count5 | 0x00000000 | 0x0003A980 |
| count6 | 0x00000064 | 0x0003A91C |
| count7 | 0x00000064 | 0x0003A91C |
| count8 | 0x0003A980 | 0x0003A980 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 FF

11111111 over 320000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000064 | 0x0004E19C |  |  |  |
| count1 | 0x00000064 | 0x0004E19C |  |  |  |
| count2 | 0x00000064 | 0x0004E19C |  |  |  |
| count3 | 0x00000064 | 0x0004E19C |  |  |  |
| count4 | 0x00000064 | 0x0004E19C |  |  |  |
| count5 | 0x00000064 | 0x0004E19C |  |  |  |
| count6 | 0x00000064 | 0x0004E19C |  |  |  |
| count7 | 0x00000064 | 0x0004E19C |  |  |  |
| count8 Validity | 0x0004E200 | 0x0004E200 |  |  |  |
| count0 | 0x00000064 | 0x0004E19C |  |  |  |
| count1 | 0x00000064 | 0x0004E19C |  |  |  |
| count2 | 0x00000064 | 0x0004E19C |  |  |  |
| count3 | 0x00000064 | 0x0004E19C |  |  |  |
| count4 | 0x00000064 | 0x0004E19C |  |  |  |
| count5 | 0x00000064 | 0x0004E19C |  |  |  |
| count6 | 0x00000064 | 0x0004E19C |  |  |  |
| count7 | 0x00000064 | 0x0004E19C |  |  |  |
| count8 Validity | 0x0004E200 | 0x0004E200 |  |  |  |
|  |  |  |  | PASS | FAIL |

## M16 1_2

Tests using the 1_2  1 from 2 CRM configuration have a more limited set of wave forms i.e. all ones all zeros or 0101 and 1010.

We specify 3(+1) frames as before.  So that the total number of samples in the 1_2 case is 2500*8*4*2 = 1600000 samples.  The modulo 3200 counter loops 1600000/3200 = 50 times.  So giving a count of 50 for a single count per loop

With the 16meg 1_2 condition we can apply the following a CRM table patterns:-

Run “C:>PCAL16  \PCALDATA\TABT000.DAT 16”     for TABT000.

Check that the results agree with the TABT000.DAT column of the following tables.

Then run “C:>PCAL16  \PCALDATA\TABT003.DAT 16”     for TABT003.

Check that the results agree with the TABT003.DAT column of the following tables.

Results are stored in 16**.RSP where ** refers to the CRM configuration e.g. FF, 53

### 16meg 1_2 FF

11111111 over 160000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000032 | 0x000270CE |  |  |  |
| count1 | 0x00000032 | 0x000270CE |  |  |  |
| count2 | 0x00000032 | 0x000270CE |  |  |  |
| count3 | 0x00000032 | 0x000270CE |  |  |  |
| count4 | 0x00000032 | 0x000270CE |  |  |  |
| count5 | 0x00000032 | 0x000270CE |  |  |  |
| count6 | 0x00000032 | 0x000270CE |  |  |  |
| count7 | 0x00000032 | 0x000270CE |  |  |  |
| count8 | 0x00027100 | 0x00027100 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 16meg 1_2 00

00000000 over 160000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000000 | 0x00000000 |  |  |  |
| count1 | 0x00000000 | 0x00000000 |  |  |  |
| count2 | 0x00000000 | 0x00000000 |  |  |  |
| count3 | 0x00000000 | 0x00000000 |  |  |  |
| count4 | 0x00000000 | 0x00000000 |  |  |  |
| count5 | 0x00000000 | 0x00000000 |  |  |  |
| count6 | 0x00000000 | 0x00000000 |  |  |  |
| count7 | 0x00000000 | 0x00000000 |  |  |  |
| count8 | 0x00000000 | 0x00000000 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 16meg 1_2 11

01010101 over 160000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000000 | 0x00013880 |  |  |  |
| count1 | 0x00000032 | 0x0001384E |  |  |  |
| count2 | 0x00000000 | 0x00013880 |  |  |  |
| count3 | 0x00000032 | 0x0001384E |  |  |  |
| count4 | 0x00000000 | 0x00013880 |  |  |  |
| count5 | 0x00000032 | 0x0001384E |  |  |  |
| count6 | 0x00000000 | 0x00013880 |  |  |  |
| count7 | 0x00000032 | 0x0001384E |  |  |  |
| count8 | 0x00013880 | 0x00013880 |  |  |  |
|  |  |  |  | PASS | FAIL |

### 16meg 1_2 22

10101010 over 160000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |
| --- | --- | --- |
| count0 | 0x00000032 | 0x0001384E |
| count1 | 0x00000000 | 0x00013880 |
| count2 | 0x00000032 | 0x0001384E |
| count3 | 0x00000000 | 0x00013880 |
| count4 | 0x00000032 | 0x0001384E |
| count5 | 0x00000000 | 0x00013880 |
| count6 | 0x00000032 | 0x0001384E |
| count7 | 0x00000000 | 0x00013880 |
| count8 | 0x00013880 | 0x00013880 |

| PASS | FAIL |
| --- | --- |

## M16 M08 M04 1_1

Tests using the 1_1  1 from 1 CRM configuration have a very limited set of wave forms i.e. all ones or all zeros.

We specify 3(+1) frames as before.  So that the total number of samples in the 1_1 case is 2500*8*4*1 = 80000 samples.  The modulo 3200 counter loops 80000/3200 = 25 times.  The results should be 1/4 of the count obtained for FF data using the same table data.

### 16meg 08meg 04meg 1_1 FF

11111111 over 80000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT003.DAT |  |  |  |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000019 | 0x00013867 |  |  |  |
| count1 | 0x00000019 | 0x00013867 |  |  |  |
| count2 | 0x00000019 | 0x00013867 |  |  |  |
| count3 | 0x00000019 | 0x00013867 |  |  |  |
| count4 | 0x00000019 | 0x00013867 |  |  |  |
| count5 | 0x00000019 | 0x00013867 |  |  |  |
| count6 | 0x00000019 | 0x00013867 |  |  |  |
| count7 | 0x00000019 | 0x00013867 |  |  |  |
| count8 | 0x00013880 | 0x00013880 |  |  |  |
|  |  |  |  | PASS | FAIL |

## M08 M04 M02 2_1

Tests using the 2_1  2 channels  from track CRM configuration also have a very  limited set of wave forms i.e. all ones or all zeros.  But they do present an interesting case.

We specify 3(+1) frames as before.  So that the total number of samples in the 1_2 case is 2500*8*4*1/2 = 40000 samples.  The modulo 3200 counter loops 40000/3200 = 12.5 times.

We use the fact that the modulo 3200 counter does not complete an integral number of loops to prove that the counting system starts and finishes at the correct point.  We set up the SRAM table TABT002.DAT which positions the data to straddle half way through the 3200 locations of RAM for S&M*V = 1.  The first 4 counts enjoy the 12 plus one half loops of the modulo 3200 counter. Where as the second set of 4 counters only see 12 loops.

In directory C:\M02\2_1

Run “C:>PCAL02  \PCALDATA\ TABT000.DAT 02”.

Check that the results agree with the TABT000.DAT column of the following tables.

Run “C:>PCAL02  \PCALDATA\ TABT001.DAT 02”.

Check that the results agree with the TABT001.DAT column of the following tables.

Run “C:>PCAL02  \PCALDATA\ TABT002.DAT 02”.

Check that the results agree with the TABT002.DAT column of the following tables.

Run “C:>PCAL02  \PCALDATA\ TABT003.DAT 02”.

Check that the results agree with the TABT003.DAT column of the following tables.

Run “C:>PCAL02  \PCALDATA\ RAND.DAT 02”.

Check that the results agree with the RAND.DAT column of the following tables.

### 02meg 2_1 FF / 04meg 2_1 FF / 08meg 2_1 FF

11111111 over 40000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT001.DAT | TABT002.DAT | TABT003.DAT | RAND.DAT |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x0000000D | 0x0000000D | 0x0000000D | 0x00009C33 | 0x00004F5A |
| count1 | 0x0000000D | 0x0000000D | 0x0000000D | 0x00009C33 | 0x00004CCB |
| count2 | 0x0000000D | 0x0000000D | 0x0000000D | 0x00009C33 | 0x00004F55 |
| count3 | 0x0000000D | 0x0000000D | 0x0000000D | 0x00009C33 | 0x00004EA8 |
| count4 | 0x0000000D | 0x0000000D | 0x0000000C | 0x00009C33 | 0x00004D51 |
| count5 | 0x0000000D | 0x0000000D | 0x0000000C | 0x00009C33 | 0x00004E33 |
| count6 | 0x0000000D | 0x0000000D | 0x0000000C | 0x00009C33 | 0x00004D87 |
| count7 | 0x0000000D | 0x0000000D | 0x0000000C | 0x00009C33 | 0x00004E4F |
| count8 | 0x00009C40 | 0x00009C40 | 0x00009C40 | 0x00009C40 | 0x00009C40 |

| PASS | FAIL |
| --- | --- |

### 02meg 2_1 00 / 04meg 2_1 00 / 08meg 2_1 00

00000000 over 40000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT001.DAT | TABT002.DAT | TABT003.DAT | RAND.DAT |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 |
| count1 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 |
| count2 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 |
| count3 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 |
| count4 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 |
| count5 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 |
| count6 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 |
| count7 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 |
| count8 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 | 0x00000000 |

| PASS | FAIL |
| --- | --- |

## M04 M02 4_1

Tests using the 4_1  2 channels  from track CRM configuration also have a very  limited set of wave forms i.e. all ones or all zeros.  But they do present an interesting case.

We specify 3(+1) frames as before.  So that the total number of samples in the 1_4 case is 2500*8*4*1/4 = 20000 samples.  The modulo 3200 counter loops 20000/3200 = 6.25 times.

We use the fact that the modulo 3200 counter does not complete an integral number of loops to prove that the counting system starts and finishes at the correct point.  We set up the SRAM table TABT001.DAT which positions the data to straddle quarter way through the 3200 locations of RAM for S&M*V = 1.  The first 4 counts enjoy the 6 plus one quarter loops of the modulo 3200 counter. Where as the second set of 4 counters only see 6 loops.

In directory C:\M02\4_1

Run “C:>PCAL02  \PCALDATA\ TABT000.DAT 02”.

Check that the results agree with the TABT000.DAT column of the following tables.

Run “C:>PCAL02  \PCALDATA\ TABT001.DAT 02”.

Check that the results agree with the TABT001.DAT column of the following tables.

Run “C:>PCAL02  \PCALDATA\ TABT002.DAT 02”.

Check that the results agree with the TABT002.DAT column of the following tables.

Run “C:>PCAL02  \PCALDATA\ TABT003.DAT 02”.

Check that the results agree with the TABT003.DAT column of the following tables.

Run “C:>PCAL02  \PCALDATA\ RAND.DAT 02”.

Check that the results agree with the RAND.DAT column of the following tables.

### 02meg 4_1 FF / 04meg 4_1 FF

11111111  over 20000 samples ( 4 input frames ).

|  | TABT000.DAT | TABT001.DAT | TABT002.DAT | TABT003.DAT | RAND.DAT |
| --- | --- | --- | --- | --- | --- |
| count0 | 0x00000007 | 0x00000007 | 0x00000006 | 0x00004E19 | 0x000027B5 |
| count1 | 0x00000007 | 0x00000007 | 0x00000006 | 0x00004E19 | 0x00002665 |
| count2 | 0x00000007 | 0x00000007 | 0x00000006 | 0x00004E19 | 0x000027B6 |
| count3 | 0x00000007 | 0x00000007 | 0x00000006 | 0x00004E19 | 0x0000275B |
| count4 | 0x00000007 | 0x00000006 | 0x00000006 | 0x00004E19 | 0x000026B1 |
| count5 | 0x00000007 | 0x00000006 | 0x00000006 | 0x00004E19 | 0x0000272C |
| count6 | 0x00000007 | 0x00000006 | 0x00000006 | 0x00004E19 | 0x000026D3 |
| count7 | 0x00000007 | 0x00000006 | 0x00000006 | 0x00004E19 | 0x00002722 |
| count8 | 0x00004E20 | 0x00004E20 | 0x00004E20 | 0x00004E20 | 0x00004E20 |

| PASS | FAIL |
| --- | --- |

## Random Table data and Stress tests for large counts.

All tests out lined so far have been based on table data with intuitively predictable results over a small number of frames producing relatively small resulting counts.  To demonstrate operation over a longer period and using essentially random table data we can run the 1_4 tests but increase the requested frame count and use table data supplied by B.A.  The frame count is determined by computation assuming that the data is all ones for all frames.  The counters will over flow after 2^32 samples.  Each frame in the 1_4 case has 80000 samples so the frame limit is set for 53687 frames.  The expected results can be computed in the same fashion as for the short tests.  We can take the counts for 4 frames and multiply up by 53687/4.

In directory C:\32M\1_4

Run “C:>PCAL32  \PCALDATA\ RAND.DAT 32”    for 4 frames.

Check that the results agree with the 4 frames column of the following tables.

Run “C:>PCAL32  \PCALDATA\RAND.DAT 32”     for 53687 frames.

Check that the results agree with the 53687 frames column of the following tables.

Results are stored in 32**.RSP where ** refers to the CRM configuration e.g. FF, 53

### 32meg 1_4 00 Random Data

000000 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x00000000 | 0x00000000 |
| count1 | 0x00000000 | 0x00000000 |
| count2 | 0x00000000 | 0x00000000 |
| count3 | 0x00000000 | 0x00000000 |
| count4 | 0x00000000 | 0x00000000 |
| count5 | 0x00000000 | 0x00000000 |
| count6 | 0x00000000 | 0x00000000 |
| count7 | 0x00000000 | 0x00000000 |
| count8 | 0x00000000 | 0x00000000 |
| count0 | 0x00000000 | 0x00000000 |
| count1 | 0x00000000 | 0x00000000 |
| count2 | 0x00000000 | 0x00000000 |
| count3 | 0x00000000 | 0x00000000 |
| count4 | 0x00000000 | 0x00000000 |
| count5 | 0x00000000 | 0x00000000 |
| count6 | 0x00000000 | 0x00000000 |
| count7 | 0x00000000 | 0x00000000 |
| count8 | 0x00000000 | 0x00000000 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 53 Random Data

00010001 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x000097F4 | 0x1F1EB4DB |
| count1 | 0x000099E8 | 0x1F851B36 |
| count2 | 0x000097F4 | 0x1F1EB4DB |
| count3 | 0x00009AB0 | 0x1FAE10F4 |
| count4 | 0x00009A4C | 0x1F999615 |
| count5 | 0x0000A53C | 0x21D70679 |
| count6 | 0x000097F4 | 0x1F1EB4DB |
| count7 | 0x00009600 | 0x1EB84E80 |
| count8 | 0x00013880 | 0x3FFFF8E0 |
| count0 | 0x0000972C | 0X1EF5BF1D |
| count1 | 0x00009538 | 0x1E8F58C2 |
| count2 | 0x00009B78 | 0x1FD706B2 |
| count3 | 0x0000A0F0 | 0x20F5BEE4 |
| count4 | 0x00009F60 | 0x20A3D368 |
| count5 | 0x00009920 | 0x1F5C2578 |
| count6 | 0x000098BC | 0x1F47AA99 |
| count7 | 0x0000A08C | 0x20E14405 |
| count8 | 0x00013880 | 0x3FFFF8E0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 5E Random Data

01000100 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x0000972C | 0X1EF5BF1D |
| count1 | 0x00009538 | 0x1E8F58C2 |
| count2 | 0x00009B78 | 0x1FD706B2 |
| count3 | 0x0000A0F0 | 0x20F5BEE4 |
| count4 | 0x00009F60 | 0x20A3D368 |
| count5 | 0x00009920 | 0x1F5C2578 |
| count6 | 0x000098BC | 0x1F47AA99 |
| count7 | 0x0000A08C | 0x20E14405 |
| count8 | 0x00013880 | 0x3FFFF8E0 |
| count0 | 0x000097F4 | 0x1F1EB4DB |
| count1 | 0x000099E8 | 0x1F851B36 |
| count2 | 0x000097F4 | 0x1F1EB4DB |
| count3 | 0x00009AB0 | 0x1FAE10F4 |
| count4 | 0x00009A4C | 0x1F999615 |
| count5 | 0x0000A53C | 0x21D70679 |
| count6 | 0x000097F4 | 0x1F1EB4DB |
| count7 | 0x00009600 | 0x1EB84E80 |
| count8 | 0x00013880 | 0x3FFFF8E0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 5F Random Data

01010101 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x00012F20 | 0x3E1473F8 |
| count1 | 0x00012F20 | 0x3E1473F8 |
| count2 | 0x0001336C | 0x3EF5BB8D |
| count3 | 0x00013BA0 | 0x40A3CFD8 |
| count4 | 0x000139AC | 0x403D697D |
| count5 | 0x00013E5C | 0x41332BF1 |
| count6 | 0x000130B0 | 0x3E665F74 |
| count7 | 0x0001368C | 0x3F999285 |
| count8 | 0x00027100 | 0x7FFFF1C0 |
| count0 | 0x00012F20 | 0x3E1473F8 |
| count1 | 0x00012F20 | 0x3E1473F8 |
| count2 | 0x0001336C | 0x3EF5BB8D |
| count3 | 0x00013BA0 | 0x40A3CFD8 |
| count4 | 0x000139AC | 0x403D697D |
| count5 | 0x00013E5C | 0x41332BF1 |
| count6 | 0x000130B0 | 0x3E665F74 |
| count7 | 0x0001368C | 0x3F999285 |
| count8 | 0x00027100 | 0x7FFFF1C0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 6F Random Data

01100110 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x00013628 | 0x3F8517A6 |
| count1 | 0x0001336C | 0x3EF5BB8D |
| count2 | 0x00013DF8 | 0x411EB112 |
| count3 | 0x00013A74 | 0x40665F3B |
| count4 | 0x00013AD8 | 0x407ADA1A |
| count5 | 0x00013114 | 0x3E7ADA53 |
| count6 | 0x00013498 | 0x3F332C2A |
| count7 | 0x00013F88 | 0x41709C8E |
| count8 | 0x00027100 | 0x7FFFF1C0 |
| count0 | 0x0001449C | 0x427AD9E1 |
| count1 | 0x0001336C | 0x3EF5BB8D |
| count2 | 0x00013CCC | 0x40E14075 |
| count3 | 0x00013BA0 | 0x40A3CFD8 |
| count4 | 0x00012F84 | 0x3E28EED7 |
| count5 | 0x000140B4 | 0x41AE0D2B |
| count6 | 0x000137B8 | 0x3FD70322 |
| count7 | 0x00013308 | 0x3EE140AE |
| count8 | 0x00027100 | 0x7FFFF1C0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 7F Random Data

01110111 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x0001CE1C | 0x5EA3CC81 |
| count1 | 0x0001CD54 | 0x5E7AD6C3 |
| count2 | 0x0001D5EC | 0x603D65ED |
| count3 | 0x0001D524 | 0x6014702F |
| count4 | 0x0001D524 | 0x6014702F |
| count5 | 0x0001D650 | 0x6051E0CC |
| count6 | 0x0001CC8C | 0x5E51E105 |
| count7 | 0x0001D588 | 0x6028EB0E |
| count8 | 0x0003A980 | 0xBFFFEAA0 |
| count0 | 0x0001DBC8 | 0x617098FE |
| count1 | 0x0001C8A4 | 0x5D85144F |
| count2 | 0x0001D844 | 0x60B84727 |
| count3 | 0x0001DC90 | 0x61998EBC |
| count4 | 0x0001CEE4 | 0x5ECCC23F |
| count5 | 0x0001D9D4 | 0x610A32A3 |
| count6 | 0x0001D074 | 0x5F1EADBB |
| count7 | 0x0001D394 | 0x5FC284B3 |
| count8 | 0x0003A980 | 0xBFFFEAA0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 A6 Random Data

00100010 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x00009EFC | 0x208F5889 |
| count1 | 0x00009E34 | 0x206662CB |
| count2 | 0x0000A280 | 0x2147AA60 |
| count3 | 0x00009984 | 0x1F70A057 |
| count4 | 0x00009B78 | 0x1FD706B2 |
| count5 | 0x000097F4 | 0x1F1EB4DB |
| count6 | 0x00009BDC | 0x1FEB8191 |
| count7 | 0x00009EFC | 0x208F5889 |
| count8 | 0x00013880 | 0x3FFFF8E0 |
| count0 | 0x0000ACA8 | 0x235C2506 |
| count1 | 0x00009984 | 0x1F70A057 |
| count2 | 0x0000A4D8 | 0x21C28B9A |
| count3 | 0x0000A0F0 | 0x20F5BEE4 |
| count4 | 0x00009538 | 0x1E8F58C2 |
| count5 | 0x00009B78 | 0x1FD706B2 |
| count6 | 0x00009FC4 | 0x20B84E47 |
| count7 | 0x00009D08 | 0x2028F22E |
| count8 | 0x00013880 | 0x3FFFF8E0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 AD Random Data

10001000 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x0000ACA8 | 0x235C2506 |
| count1 | 0x00009984 | 0x1F70A057 |
| count2 | 0x0000A4D8 | 0x21C28B9A |
| count3 | 0x0000A0F0 | 0x20F5BEE4 |
| count4 | 0x00009538 | 0x1E8F58C2 |
| count5 | 0x00009B78 | 0x1FD706B2 |
| count6 | 0x00009FC4 | 0x20B84E47 |
| count7 | 0x00009D08 | 0x2028F22E |
| count8 | 0x00013880 | 0x3FFFF8E0 |
| count0 | 0x00009EFC | 0x208F5889 |
| count1 | 0x00009E34 | 0x206662CB |
| count2 | 0x0000A280 | 0x2147AA60 |
| count3 | 0x00009984 | 0x1F70A057 |
| count4 | 0x00009B78 | 0x1FD706B2 |
| count5 | 0x000097F4 | 0x1F1EB4DB |
| count6 | 0x00009BDC | 0x1FEB8191 |
| count7 | 0x00009EFC | 0x208F5889 |
| count8 | 0x00013880 | 0x3FFFF8E0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 AF Random Data

10011001 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x00014BA4 | 0x43EB7D8F |
| count1 | 0x000137B8 | 0x3FD70322 |
| count2 | 0x00014758 | 0x430A35FA |
| count3 | 0x00013A74 | 0x40665F3B |
| count4 | 0x000130B0 | 0x3E665F74 |
| count5 | 0x0001336C | 0x3EF5BB8D |
| count6 | 0x00013BA0 | 0x40A3CFD8 |
| count7 | 0x00013C04 | 0x40B84AB7 |
| count8 | 0x00027100 | 0x7FFFF1C0 |
| count0 | 0x00014BA4 | 0x43EB7D8F |
| count1 | 0x000137B8 | 0x3FD70322 |
| count2 | 0x00014758 | 0x430A35FA |
| count3 | 0x00013A74 | 0x40665F3B |
| count4 | 0x000130B0 | 0x3E665F74 |
| count5 | 0x0001336C | 0x3EF5BB8D |
| count6 | 0x00013BA0 | 0x40A3CFD8 |
| count7 | 0x00013C04 | 0x40B84AB7 |
| count8 | 0x00027100 | 0x7FFFF1C0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 B7 Random Data

00110011 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x000136F0 | 0x3FAE0D64 |
| count1 | 0x0001381C | 0x3FEB7E01 |
| count2 | 0x00013A74 | 0x40665F3B |
| count3 | 0x00013434 | 0x3F1EB14B |
| count4 | 0x000135C4 | 0x3F709CC7 |
| count5 | 0x00013D30 | 0x40F5BB54 |
| count6 | 0x000133D0 | 0x3F0A366C |
| count7 | 0x000134FC | 0x3F47A709 |
| count8 | 0x00027100 | 0x7FFFF1C0 |
| count0 | 0x000143D4 | 0x4251E423 |
| count1 | 0x00012EBC | 0x3DFFF919 |
| count2 | 0x00014050 | 0x4199924C |
| count3 | 0x000141E0 | 0x41EB7DC8 |
| count4 | 0x00013498 | 0x3F332C2A |
| count5 | 0x00013498 | 0x3F332C2A |
| count6 | 0x00013880 | 0x3FFFF8E0 |
| count7 | 0x00013D94 | 0x410A3633 |
| count8 | 0x00027100 | 0x7FFFF1C0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 BF Random Data

10111011 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x0001E398 | 0x630A326A |
| count1 | 0x0001D1A0 | 0x5F5C1E58 |
| count2 | 0x0001DF4C | 0x6228EAD5 |
| count3 | 0x0001D524 | 0x6014702F |
| count4 | 0x0001CAFC | 0x5DFFF589 |
| count5 | 0x0001D8A8 | 0x60CCC206 |
| count6 | 0x0001D394 | 0x5FC284B3 |
| count7 | 0x0001D204 | 0x5F709937 |
| count8 | 0x0003A980 | 0xBFFFEAA0 |
| count0 | 0x0001E2D0 | 0x62E13CAC |
| count1 | 0x0001CCF0 | 0x5E665BE4 |
| count2 | 0x0001E2D0 | 0x62E13CAC |
| count3 | 0x0001DB64 | 0x615C1E1F |
| count4 | 0x0001D010 | 0x5F0A32DC |
| count5 | 0x0001CC8C | 0x5E51E105 |
| count6 | 0x0001D45C | 0x5FEB7A71 |
| count7 | 0x0001DC90 | 0x61998EBC |
| count8 | 0x0003A980 | 0xBFFFEAA0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 DB Random Data

10011001 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x0001449C | 0x427AD9E1 |
| count1 | 0x0001336C | 0x3EF5BB8D |
| count2 | 0x00013CCC | 0x40E14075 |
| count3 | 0x00013BA0 | 0x40A3CFD8 |
| count4 | 0x00012F84 | 0x3E28EED7 |
| count5 | 0x000140B4 | 0x41AE0D2B |
| count6 | 0x000137B8 | 0x3FD70322 |
| count7 | 0x00013308 | 0x3EE140AE |
| count8 | 0x00027100 | 0x7FFFF1C0 |
| count0 | 0x00013628 | 0x3F8517A6 |
| count1 | 0x0001336C | 0x3EF5BB8D |
| count2 | 0x00013DF8 | 0x411EB112 |
| count3 | 0x00013A74 | 0x40665F3B |
| count4 | 0x00013AD8 | 0x407ADA1A |
| count5 | 0x00013114 | 0x3E7ADA53 |
| count6 | 0x00013498 | 0x3F332C2A |
| count7 | 0x00013F88 | 0x41709C8E |
| count8 | 0x00027100 | 0x7FFFF1C0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 DE Random Data

11001100 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x000143D4 | 0x4251E423 |
| count1 | 0x00012EBC | 0x3DFFF919 |
| count2 | 0x00014050 | 0x4199924C |
| count3 | 0x000141E0 | 0x41EB7DC8 |
| count4 | 0x00013498 | 0x3F332C2A |
| count5 | 0x00013498 | 0x3F332C2A |
| count6 | 0x00013880 | 0x3FFFF8E0 |
| count7 | 0x00013D94 | 0x410A3633 |
| count8 | 0x00027100 | 0x7FFFF1C0 |
| count0 | 0x000136F0 | 0x3FAE0D64 |
| count1 | 0x0001381C | 0x3FEB7E01 |
| count2 | 0x00013A74 | 0x40665F3B |
| count3 | 0x00013434 | 0x3F1EB14B |
| count4 | 0x000135C4 | 0x3F709CC7 |
| count5 | 0x00013D30 | 0x40F5BB54 |
| count6 | 0x000133D0 | 0x3F0A366C |
| count7 | 0x000134FC | 0x3F47A709 |
| count8 | 0x00027100 | 0x7FFFF1C0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 DF Random Data

11011101 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x0001DBC8 | 0x617098FE |
| count1 | 0x0001C8A4 | 0x5D85144F |
| count2 | 0x0001D844 | 0x60B84727 |
| count3 | 0x0001DC90 | 0x61998EBC |
| count4 | 0x0001CEE4 | 0x5ECCC23F |
| count5 | 0x0001D9D4 | 0x610A32A3 |
| count6 | 0x0001D074 | 0x5F1EADBB |
| count7 | 0x0001D394 | 0x5FC284B3 |
| count8 | 0x0003A980 | 0xBFFFEAA0 |
| count0 | 0x0001CE1C | 0x5EA3CC81 |
| count1 | 0x0001CD54 | 0x5E7AD6C3 |
| count2 | 0x0001D5EC | 0x603D65ED |
| count3 | 0x0001D524 | 0x6014702F |
| count4 | 0x0001D524 | 0x6014702F |
| count5 | 0x0001D650 | 0x6051E0CC |
| count6 | 0x0001CC8C | 0x5E51E105 |
| count7 | 0x0001D588 | 0x6028EB0E |
| count8 | 0x0003A980 | 0xBFFFEAA0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 EF Random Data

11101110 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x0001E2D0 | 0x62E13CAC |
| count1 | 0x0001CCF0 | 0x5E665BE4 |
| count2 | 0x0001E2D0 | 0x62E13CAC |
| count3 | 0x0001DB64 | 0x615C1E1F |
| count4 | 0x0001D010 | 0x5F0A32DC |
| count5 | 0x0001CC8C | 0x5E51E105 |
| count6 | 0x0001D45C | 0x5FEB7A71 |
| count7 | 0x0001DC90 | 0x61998EBC |
| count8 | 0x0003A980 | 0xBFFFEAA0 |
| count0 | 0x0001E398 | 0x630A326A |
| count1 | 0x0001D1A0 | 0x5F5C1E58 |
| count2 | 0x0001DF4C | 0x6228EAD5 |
| count3 | 0x0001D524 | 0x6014702F |
| count4 | 0x0001CAFC | 0x5DFFF589 |
| count5 | 0x0001D8A8 | 0x60CCC206 |
| count6 | 0x0001D394 | 0x5FC284B3 |
| count7 | 0x0001D204 | 0x5F709937 |
| count8 | 0x0003A980 | 0xBFFFEAA0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 FF Random Data

11111111 Over 4 frames and 53687 frames

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x00027AC4 | 0x81FFF187 |
| count1 | 0x000266D8 | 0x7DEB771A |
| count2 | 0x00027AC4 | 0x81FFF187 |
| count3 | 0x00027614 | 0x810A2F13 |
| count4 | 0x00026A5C | 0x7EA3C8F1 |
| count5 | 0x000271C8 | 0x8028E77E |
| count6 | 0x00026C50 | 0x7F0A2F4C |
| count7 | 0x00027290 | 0x8051DD3C |
| count8 | 0x0004E200 | 0xFFFFE380 |
| count0 | 0x00027AC4 | 0x81FFF187 |
| count1 | 0x000266D8 | 0x7DEB771A |
| count2 | 0x00027AC4 | 0x81FFF187 |
| count3 | 0x00027614 | 0x810A2F13 |
| count4 | 0x00026A5C | 0x7EA3C8F1 |
| count5 | 0x000271C8 | 0x8028E77E |
| count6 | 0x00026C50 | 0x7F0A2F4C |
| count7 | 0x00027290 | 0x8051DD3C |
| count8 | 0x0004E200 | 0xFFFFE380 |

| PASS | FAIL |
| --- | --- |

## Tests for non zero data in the validity equals zero table section.

Tests so far have been based on table data with all zeros in the lower area corresponding to validity equals zero  These tests fill all the lower section with a non zero table and fill the upper section with zeroes.  Then we perform tests in the normal way.

Run “C:>PCALZ32  \PCALDATA\ TABT005.DAT 32”    for 4 frames.

Check that the results agree with the 4 frames column of the following tables.

Run “C:>PCALZ32 \PCALDATA\TABT005.DAT 32”     for 53687 frames.

Check that the results agree with the 53687 frames column of the following tables.

Results are stored in 32**.RSP where ** refers to the CRM configuration e.g. FF, 53

### 32meg 1_4 00 Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file RANDZ.DAT and TABT005.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x0002709C | 0x7FEB76E1 |
| count1 | 0x000275B0 | 0x80F5B434 |
| count2 | 0x000267A0 | 0x7E146CD8 |
| count3 | 0x00027B8C | 0x8228E745 |
| count4 | 0x00026A5C | 0x7EA3C8F1 |
| count5 | 0x00026D18 | 0x7F33250A |
| count6 | 0x000267A0 | 0x7E146CD8 |
| count7 | 0x00027F10 | 0x82E1391C |
| count8 | 0x00000000 | 0x00000000 |
| count0. | 0x0002709C | 0x7FEB76E1 |
| count1. | 0x000275B0 | 0x80F5B434 |
| count2. | 0x000267A0 | 0x7E146CD8 |
| count3. | 0x00027B8C | 0x8228E745 |
| count4. | 0x00026A5C | 0x7EA3C8F1 |
| count5. | 0x00026D18 | 0x7F33250A |
| count6. | 0x000267A0 | 0x7E146CD8 |
| count7. | 0x00027F10 | 0x82E1391C |
| count8. | 0x00000000 | 0x00000000 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 53 Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file RANDZ.DAT and TABT005.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x0001E3FC | 0x631EAD49 |
| count1 | 0x0001D9D4 | 0x610A32A3 |
| count2 | 0x0001C3F4 | 0x5C8F51DB |
| count3 | 0x0001DC2C | 0x618513DD |
| count4 | 0x0001DE20 | 0x61EB7A38 |
| count5 | 0x0001ECF8 | 0x64F5B752 |
| count6 | 0x0001C64C | 0x5D0A3315 |
| count7 | 0x0001D394 | 0x5FC284B3 |
| count8 | 0x00013880 | 0x3FFFF8E0 |
| count0 | 0x0001C584 | 0x5CE13D57 |
| count1 | 0x0001DE84 | 0x61FFF517 |
| count2 | 0x0001DEE8 | 0x62146FF6 |
| count3 | 0x0001C138 | 0x5BFFF5C2 |
| count4 | 0x0001D45C | 0x5FEB7A71 |
| count5 | 0x0001C520 | 0x5CCCC278 |
| count6 | 0x0001CB60 | 0x5E147068 |
| count7 | 0x0001DB00 | 0x6147A340 |
| count8 | 0x00013880 | 0x3FFFF8E0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 5E Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file RANDZ.DAT and TABT005.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x0001C3F4 | 0x5C8F51DB |
| count1 | 0x0001DA9C | 0x61332861 |
| count2 | 0x0001D4C0 | 0x5FFFF550 |
| count3 | 0x0001DB00 | 0x6147A340 |
| count4 | 0x0001CA34 | 0x5DD6FFCB |
| count5 | 0x0001D970 | 0x60F5B7C4 |
| count6 | 0x0001CA98 | 0x5DEB7AAA |
| count7 | 0x0001DC2C | 0x618513DD |
| count8 | 0x00013880 | 0x3FFFF8E0 |
| count0 | 0x0001E014 | 0x6251E093 |
| count1 | 0x0001DC2C | 0x618513DD |
| count2 | 0x0001E460 | 0x63332828 |
| count3 | 0x0001C714 | 0x5D3328D3 |
| count4 | 0x0001CC8C | 0x5E51E105 |
| count5 | 0x0001CBC4 | 0x5E28EB47 |
| count6 | 0x0001CA34 | 0x5DD6FFCB |
| count7 | 0x0001C778 | 0x5D47A3B2 |
| count8 | 0x00013880 | 0x3FFFF8E0 |

| PASS | FAIL |
| --- | --- |

### 2meg 1_4 5F Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file RANDZ.DAT and TABT005.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x00012D90 | 0x3DC2887C |
| count1 | 0x00013B3C | 0x408F54F9 |
| count2 | 0x0001381C | 0x3FEB7E01 |
| count3 | 0x0001417C | 0x41D702E9 |
| count4 | 0x000131DC | 0x3EA3D011 |
| count5 | 0x00013EC0 | 0x4147A6D0 |
| count6 | 0x00013308 | 0x3EE140AE |
| count7 | 0x00013628 | 0x3F8517A6 |
| count8 | 0x00027100 | 0x7FFFF1C0 |
| count0 | 0x00013CCC | 0x40E14075 |
| count1 | 0x00013FEC | 0x4185176D |
| count2 | 0x000143D4 | 0x4251E423 |
| count3 | 0x00012EBC | 0x3DFFF919 |
| count4 | 0x00013BA0 | 0x40A3CFD8 |
| count5 | 0x00012D90 | 0x3DC2887C |
| count6 | 0x00012A0C | 0x3D0A36A5 |
| count7 | 0x00013240 | 0x3EB84AF0 |
| count8 | 0x00027100 | 0x7FFFF1C0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 6F Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file RANDZ.DAT and TABT005.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x00012B38 | 0x3D47A742 |
| count1 | 0x00013FEC | 0x4185176D |
| count2 | 0x00013F88 | 0x41709C8E |
| count3 | 0x00013A10 | 0x4051E45C |
| count4 | 0x00013308 | 0x3EE140AE |
| count5 | 0x0001368C | 0x3F999285 |
| count6 | 0x00013D94 | 0x410A3633 |
| count7 | 0x00013628 | 0x3F8517A6 |
| count8 | 0x00027100 | 0x7FFFF1C0 |
| count0 | 0x00014884 | 0x4347A697 |
| count1 | 0x00014564 | 0x42A3CF9F |
| count2 | 0x00013948 | 0x4028EE9E |
| count3 | 0x00013E5C | 0x41332BF1 |
| count4 | 0x00013628 | 0x3F8517A6 |
| count5 | 0x000146F4 | 0x42F5BB1B |
| count6 | 0x000130B0 | 0x3E665F74 |
| count7 | 0x000130B0 | 0x3E665F74 |
| count8 | 0x00027100 | 0x7FFFF1C0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 7F Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file TABZ001.DAT and TABT000.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x00009150 | 0x1DC28C0C |
| count1 | 0x00009D6C | 0x203D6D0D |
| count2 | 0x00009FC4 | 0x20B84E47 |
| count3 | 0x00009EFC | 0x208F5889 |
| count4 | 0x000091B4 | 0x1DD706EB |
| count5 | 0x00009C40 | 0x1FFFFC70 |
| count6 | 0x0000A154 | 0x210A39C3 |
| count7 | 0x00009A4C | 0x1F999615 |
| count8 | 0x0003A980 | 0xBFFFEAA0 |
| count0 | 0x0000A348 | 0x2170A01E |
| count1 | 0x0000A280 | 0x2147AA60 |
| count2 | 0x0000A280 | 0x2147AA60 |
| count3 | 0x0000A028 | 0x20CCC926 |
| count4 | 0x00009BDC | 0x1FEB8191 |
| count5 | 0x000099E8 | 0x1F851B36 |
| count6 | 0x000091B4 | 0x1DD706EB |
| count7 | 0x000097F4 | 0x1F1EB4DB |
| count8 | 0x0003A980 | 0xBFFFEAA0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 A6 Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file TABZ001.DAT and TABT000.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x0001D6B4 | 0x60665BAB |
| count1 | 0x0001D588 | 0x6028EB0E |
| count2 | 0x0001E140 | 0x628F5130 |
| count3 | 0x0001CCF0 | 0x5E665BE4 |
| count4 | 0x0001C138 | 0x5BFFF5C2 |
| count5 | 0x0001CDB8 | 0x5E8F51A2 |
| count6 | 0x0001DC2C | 0x618513DD |
| count7 | 0x0001D010 | 0x5F0A32DC |
| count8 | 0x00013880 | 0x3FFFF8E0 |
| count0 | 0x0001D4C0 | 0x5FFFF550 |
| count1 | 0x0001DF4C | 0x6228EAD5 |
| count2 | 0x0001CA98 | 0x5DEB7AAA |
| count3 | 0x0001DE84 | 0x61FFF517 |
| count4 | 0x0001DB64 | 0x615C1E1F |
| count5 | 0x0001D9D4 | 0x610A32A3 |
| count6 | 0x0001C0D4 | 0x5BEB7AE3 |
| count7 | 0x0001DC90 | 0x61998EBC |
| count8 | 0x00013880 | 0x3FFFF8E0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 AD Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file TABZ001.DAT and TABT000.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x0001E078 | 0x62665B72 |
| count1 | 0x0001E140 | 0x628F5130 |
| count2 | 0x0001E2D0 | 0x62E13CAC |
| count3 | 0x0001D394 | 0x5FC284B3 |
| count4 | 0x0001CE1C | 0x5EA3CC81 |
| count5 | 0x0001D268 | 0x5F851416 |
| count6 | 0x0001CE1C | 0x5EA3CC81 |
| count7 | 0x0001C908 | 0x5D998F2E |
| count8 | 0x00013880 | 0x3FFFF8E0 |
| count0 | 0x0001CCF0 | 0x5E665BE4 |
| count1 | 0x0001DFB0 | 0x623D65B4 |
| count2 | 0x0001D330 | 0x5FAE09D4 |
| count3 | 0x0001D6B4 | 0x60665BAB |
| count4 | 0x0001C458 | 0x5CA3CCBA |
| count5 | 0x0001E3FC | 0x631EAD49 |
| count6 | 0x0001D7E0 | 0x60A3CC48 |
| count7 | 0x0001D90C | 0x60E13CE5 |
| count8 | 0x00013880 | 0x3FFFF8E0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 AF Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file TABZ001.DAT and TABT000.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x00013D30 | 0x40F5BB54 |
| count1 | 0x00013EC0 | 0x4147A6D0 |
| count2 | 0x00014050 | 0x4199924C |
| count3 | 0x0001336C | 0x3EF5BB8D |
| count4 | 0x00013240 | 0x3EB84AF0 |
| count5 | 0x00013880 | 0x3FFFF8E0 |
| count6 | 0x00013C68 | 0x40CCC596 |
| count7 | 0x00013114 | 0x3E7ADA53 |
| count8 | 0x00027100 | 0x7FFFF1C0 |
| count0 | 0x00013BA0 | 0x40A3CFD8 |
| count1 | 0x00014244 | 0x41FFF8A7 |
| count2 | 0x0001336C | 0x3EF5BB8D |
| count3 | 0x000137B8 | 0x3FD70322 |
| count4 | 0x000132A4 | 0x3ECCC5CF |
| count5 | 0x000147BC | 0x431EB0D9 |
| count6 | 0x0001368C | 0x3F999285 |
| count7 | 0x00013EC0 | 0x4147A6D0 |
| count8 | 0x00027100 | 0x7FFFF1C0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 B7 Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file RANDZ.DAT and TABT005.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x00013CCC | 0x40E14075 |
| count1 | 0x000139AC | 0x403D697D |
| count2 | 0x000141E0 | 0x41EB7DC8 |
| count3 | 0x000126EC | 0x3C665FAD |
| count4 | 0x000130B0 | 0x3E665F74 |
| count5 | 0x000131DC | 0x3EA3D011 |
| count6 | 0x00013880 | 0x3FFFF8E0 |
| count7 | 0x00012F84 | 0x3E28EED7 |
| count8 | 0x00027100 | 0x7FFFF1C0 |
| count0 | 0x000132A4 | 0x3ECCC5CF |
| count1 | 0x00013D30 | 0x40F5BB54 |
| count2 | 0x000134FC | 0x3F47A709 |
| count3 | 0x00013C04 | 0x40B84AB7 |
| count4 | 0x00013880 | 0x3FFFF8E0 |
| count5 | 0x00013D30 | 0x40F5BB54 |
| count6 | 0x00012944 | 0x3CE140E7 |
| count7 | 0x000141E0 | 0x41EB7DC8 |
| count8 | 0x00027100 | 0x7FFFF1C0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 BF Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file RANDZ.DAT and TABT005.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x0000A348 | 0x2170A01E |
| count1 | 0x00009C40 | 0x1FFFFC70 |
| count2 | 0x0000A08C | 0x20E14405 |
| count3 | 0x00009858 | 0x1F332FBA |
| count4 | 0x000090EC | 0x1DAE112D |
| count5 | 0x00009E34 | 0x206662CB |
| count6 | 0x0000A028 | 0x20CCC926 |
| count7 | 0x00009538 | 0x1E8F58C2 |
| count8 | 0x0003A980 | 0xBFFFEAA0 |
| count0 | 0x00009664 | 0x1ECCC95F |
| count1 | 0x00009F60 | 0x20A3D368 |
| count2 | 0x00009CA4 | 0x2014774F |
| count3 | 0x00009984 | 0x1F70A057 |
| count4 | 0x00009858 | 0x1F332FBA |
| count5 | 0x00009AB0 | 0x1FAE10F4 |
| count6 | 0x00009790 | 0x1F0A39FC |
| count7 | 0x0000A604 | 0x21FFFC37 |
| count8 | 0x0003A980 | 0xBFFFEAA0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 DB Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file RANDZ.DAT and TABT005.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x0001417C | 0x41D702E9 |
| count1 | 0x000140B4 | 0x41AE0D2B |
| count2 | 0x00012F20 | 0x3E1473F8 |
| count3 | 0x000140B4 | 0x41AE0D2B |
| count4 | 0x00013A74 | 0x40665F3B |
| count5 | 0x00014F8C | 0x44B84A45 |
| count6 | 0x000130B0 | 0x3E665F74 |
| count7 | 0x00013498 | 0x3F332C2A |
| count8 | 0x00027100 | 0x7FFFF1C0 |
| count0 | 0x0001336C | 0x3EF5BB8D |
| count1 | 0x00013FEC | 0x4185176D |
| count2 | 0x00014118 | 0x41C2880A |
| count3 | 0x000129A8 | 0x3CF5BBC6 |
| count4 | 0x00014118 | 0x41C2880A |
| count5 | 0x00012DF4 | 0x3DD7035B |
| count6 | 0x00013498 | 0x3F332C2A |
| count7 | 0x00013628 | 0x3F8517A6 |
| count8 | 0x00027100 | 0x7FFFF1C0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 DE Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file RANDZ.DAT and TABT005.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x000132A4 | 0x3ECCC5CF |
| count1 | 0x00013D30 | 0x40F5BB54 |
| count2 | 0x000134FC | 0x3F47A709 |
| count3 | 0x00013C04 | 0x40B84AB7 |
| count4 | 0x00013880 | 0x3FFFF8E0 |
| count5 | 0x00013D30 | 0x40F5BB54 |
| count6 | 0x00012944 | 0x3CE140E7 |
| count7 | 0x000141E0 | 0x41EB7DC8 |
| count8 | 0x00027100 | 0x7FFFF1C0 |
| count0 | 0x00013CCC | 0x40E14075 |
| count1 | 0x000139AC | 0x403D697D |
| count2 | 0x000141E0 | 0x41EB7DC8 |
| count3 | 0x000126EC | 0x3C665FAD |
| count4 | 0x000130B0 | 0x3E665F74 |
| count5 | 0x000131DC | 0x3EA3D011 |
| count6 | 0x00013880 | 0x3FFFF8E0 |
| count7 | 0x00012F84 | 0x3E28EED7 |
| count8 | 0x00027100 | 0x7FFFF1C0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 DF Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file RANDZ.DAT and TABT005.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x00009C40 | 0x1FFFFC70 |
| count1 | 0x00009DD0 | 0x2051E7EC |
| count2 | 0x00009858 | 0x1F332FBA |
| count3 | 0x0000A280 | 0x2147AA60 |
| count4 | 0x0000A028 | 0x20CCC926 |
| count5 | 0x0000A280 | 0x2147AA60 |
| count6 | 0x000091B4 | 0x1DD706EB |
| count7 | 0x00009BDC | 0x1FEB8191 |
| count8 | 0x0003A980 | 0xBFFFEAA0 |
| count0 | 0x00009984 | 0x1F70A057 |
| count1 | 0x00009D6C | 0x203D6D0D |
| count2 | 0x0000A154 | 0x210A39C3 |
| count3 | 0x00008E94 | 0x1D332FF3 |
| count4 | 0x00009FC4 | 0x20B84E47 |
| count5 | 0x000093A8 | 0x1E3D6D46 |
| count6 | 0x00009858 | 0x1F332FBA |
| count7 | 0x00009A4C | 0x1F999615 |
| count8 | 0x0003A980 | 0xBFFFEAA0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 EF Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file RANDZ.DAT and TABT005.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x000099E8 | 0x1F851B36 |
| count1 | 0x0000A280 | 0x2147AA60 |
| count2 | 0x00009FC4 | 0x20B84E47 |
| count3 | 0x00009B14 | 0x1FC28BD3 |
| count4 | 0x0000A154 | 0x210A39C3 |
| count5 | 0x00009A4C | 0x1F999615 |
| count6 | 0x00009C40 | 0x1FFFFC70 |
| count7 | 0x00009BDC | 0x1FEB8191 |
| count8 | 0x0003A980 | 0xBFFFEAA0 |
| count0 | 0x0000A53C | 0x21D70679 |
| count1 | 0x0000A2E4 | 0x215C253F |
| count2 | 0x000096C8 | 0x1EE1443E |
| count3 | 0x00009E34 | 0x206662CB |
| count4 | 0x00009A4C | 0x1F999615 |
| count5 | 0x0000AD0C | 0x23709FE5 |
| count6 | 0x00009EFC | 0x208F5889 |
| count7 | 0x000098BC | 0x1F47AA99 |
| count8 | 0x0003A980 | 0xBFFFEAA0 |

| PASS | FAIL |
| --- | --- |

### 32meg 1_4 FF Non Zero V=0 Table Data

Over 4 frames and 53687 frames Using zerodata file RANDZ.DAT and TABT005.DAT

|  | 4 FRAMES | 53687 FRAMES |
| --- | --- | --- |
| count0 | 0x00000000 | 0x00000000 |
| count1 | 0x00000000 | 0x00000000 |
| count2 | 0x00000000 | 0x00000000 |
| count3 | 0x00000000 | 0x00000000 |
| count4 | 0x00000000 | 0x00000000 |
| count5 | 0x00000000 | 0x00000000 |
| count6 | 0x00000000 | 0x00000000 |
| count7 | 0x00000000 | 0x00000000 |
| count8 | 0x0004E200 | 0xFFFFE380 |
| count0 | 0x00000000 | 0x00000000 |
| count1 | 0x00000000 | 0x00000000 |
| count2 | 0x00000000 | 0x00000000 |
| count3 | 0x00000000 | 0x00000000 |
| count4 | 0x00000000 | 0x00000000 |
| count5 | 0x00000000 | 0x00000000 |
| count6 | 0x00000000 | 0x00000000 |
| count7 | 0x00000000 | 0x00000000 |
| count8 | 0x0004E200 | 0xFFFFE380 |

| PASS | FAIL |
| --- | --- |

# 5	Notes For Test Method.

The PCM hardware tests are currently performed using the following method.

A Station Unit rack is populated with 1 SUCC, 2 DIM, 3 TRM,1 CRM ,1 PCM, 1 DCM, 2 DMM, 1 SUIM.  The SUCC code is downloaded from the server in the normal way.  Neither the SUCC code or PCM  micro code have been modified for the proper operational use.  They have only been modified to run with the new Xilinx configuration.  The testing is performed mainly at the fastest data rate 32/36MHz.

The C3 is implemented using a PC, a dedicated batch file pcaltest.bat uses SUTEST as in other tests.  Pcaltest is invoked with two parameters:-

“pcaltest pcaltabtable crmtabdir”

The parameter pcaltabable is the full path used to call up the file to load the PCM SRAM.   Initially pcaltabf.dat is used which loads all ones in the SRAM for S=M=V=1 and zeros every where else.  The parameter crmtabdir is used as part of a path name to collect the CRM configuration files.

The actual data flow from the CRM to the PCM is gated by the signal Channel Frame (CF).  This signal has a nominal duty cycle of 1 off to 8 on when the channel data has been reconstituted 1:1.  This is due to the clocking arrangements in the CRM and the editing out of the validity bit from the head of a data byte from the TRM.  The validity bit is then replicated for each of the data bits in the byte.

CF

0111111110111111110111111110111111110111111110111111110...

The main test CRM configurations are based on a 2 bit 4:1 multiplex.

When we use the 4:1 mux the CF signal looks as below.

CF

000011111111111111111111111111111111000011111111...

DATA

012301230123012301230123012301230123012301230123...

I have also included the nominal data track source 0,1,2 or 3 for the data bits.

The CRM when configured in this way can stimulate the PCM with a selection of wave forms.  By setting the crossbar switch such that the inputs 0,1,2 and 3 are fixed at logic either logic 1 or 0.

Keith has provided CRM Crossbar configurations (see the later section CRM Configurations)which give specific wave forms.  In the CRM when configured 2 bit 4:1 multiplex the validity bit is generated as the logical AND of the S and M bits so we can generate the following wave forms.

Taking the following example files, the validity bit stream which can be counted is :-

S 10101010		AF

M 11111111

V 10101010

CF

000011111111111111111111111111111111000011111111...

VALIDITY

000010101010101010101010101010101010000010101010...

S 01010101		5F

M 11111111

V 01010101

CF

000011111111111111111111111111111111000011111111...

VALIDITY

000001010101010101010101010101010101000001010101...

S 01010101		53

M 00110011

V 00010001

CF

000011111111111111111111111111111111000011111111...

VALIDITY

000000010001000100010001000100010001000000010001...

S 10101010		A8

M 10001000

V 10001000

CF

000011111111111111111111111111111111000011111111...

VALIDITY

000010001000100010001000100010001000000010001000...

Bearing in mind the accumulator hardware only counts when the validity bit is true and that the channel frame effectively places discontinuities in all SMV data, these wave forms give quite stringent tests for the correct timing of the data to be counted from the SRAM.  The A8 data tests the timing of the first data bit in a set.

I have generated some test software which simulates the action of the accumulators, it gives the expected counts for given input wave forms and SRAM table data.  There are limitations to the parameter space tested by this method, but it does, we feel, cover the known operational requirement.

By setting the SRAM table data to all zeros except in specific places we can predict what counts we should get.  e.g. for a single bit set in the SRAM we would expect to see a count of 25.  This is because the total number of bits in a frame is 2500 * 8 * 4  and the modulo 3200 counter must loop 25 times.  The stimulus data repeats effectively every four bits, so by placing the SRAM ones data at the first count of the modulo 3200 address and selecting the AF stimulus case above we would expect to count 25.  For the 5F stimulus case zero would be expected.

During testing we are currently using a RAM table with content 0x01,0x02,0x04,0x08,0x10,0x20,0x40,0x80, followed by all zeros for the S=M=V condition.

This gives registration to the single clock level.  We can use the logic analyser to verify all observable wave forms.

Due to the way in which the CRM actually implements the 4:1 mux 2 bit data mode the channel data does not appear the same on all channels.  Only the even numbered tracks are normally used. the odd channels are created from the same base tracks but in a different multiplex order.

For example channel 0 is built from following TRM tracks

SIGN 0,1,2,3

MAG 4,5,6,7

Channel 1

SIGN 6,7,4,5

MAG 2,3,0,1

In effect this means that all odd channel counts will be based on a different input wave form.  Giving two test cases per test.

# 6	PCAL Table Contents.

The SRAM table content defined in TABT000.DAT gives a single 1 bit in each of the first 8 locations in memory for SIGN,MAGNITUDE and VALIDITY equal to one.

TABT000.DAT

This file has the following bytes at the location shown:-

9649 0X01

9650 0X02

9651 0X04

9652 0X08

9653 0X10

9654 0X20

9655 0X40

9656 0X80

All other data bytes are 0x00.  This puts the data at the

first eight of 3200 locations of RAM for SIGN,MAGNITUDE,

VALIDITY equal to one.

TABT001.DAT

This file has the following bytes at the location shown:-

10445 0X01

10446 0X02

10447 0X04

10448 0X08

10449 0X10

10450 0X20

10451 0X40

10452 0X80

All other data bytes are 0x00.  This puts the data to

straddle a quarter way through the 3200 locations of RAM

for SIGN,MAGNITUDE,VALIDITY equal to one.

TABT002.DAT

This file has the following bytes at the location shown:-

11245 0x01

11246 0x02

11247 0X04

11248 0X08

11249 0X10

11250 0X20

11251 0X40

11252 0X80

All other data bytes are 0x00.  This puts the data to

straddle half way through the 3200 locations of RAM

for SIGN,MAGNITUDE,VALIDITY equal to one.

TABT003.DAT

This file has the following bytes at the location shown:-

9649 0XFE

9650 0XFD

9651 0XFB

9652 0XF7

9653 0XEF

9654 0XDF

9655 0XBF

9656 0X7F

All other data bytes are 0xFF.  This puts the data at the

first eight of 3200 locations of RAM for SIGN,MAGNITUDE,

VALIDITY equal to one.

TABT004.DAT

All data bytes are 0xFF.

TABT005.DAT

All data bytes are 0x00.

# 7	CRM Output Patterns For Testing The PCM.

Pattern 00

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 0 | 0 | 0 | 0 | 00000000 | 00000000 |
| Magnitude | 0 | 0 | 0 | 0 | 00000000 | 00000000 |
| Validity | - | - | - | - | 00000000 | 00000000 |

Pattern 53

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 0 | 1 | 0 | 1 | 01010101 | 11001100 |
| Magnitude | 0 | 0 | 1 | 1 | 00110011 | 01010101 |
| Validity | - | - | - | - | 00010001 | 01000100 |

Pattern A6

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 1 | 0 | 1 | 0 | 10101010 | 10011001 |
| Magnitude | 0 | 1 | 1 | 0 | 01100110 | 10101010 |
| Validity | - | - | - | - | 00100010 | 10001000 |

Pattern B7

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 1 | 0 | 1 | 1 | 10111011 | 11011101 |
| Magnitude | 0 | 1 | 1 | 1 | 01110111 | 11101110 |
| Validity | - | - | - | - | 00110011 | 11001100 |

Pattern 5E

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 0 | 1 | 0 | 1 | 01010101 | 10111011 |
| Magnitude | 1 | 1 | 1 | 0 | 11101110 | 01010101 |
| Validity | - | - | - | - | 01000100 | 00010001 |

Pattern 5F

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 0 | 1 | 0 | 1 | 01010101 | 11111111 |
| Magnitude | 1 | 1 | 1 | 1 | 11111111 | 01010101 |
| Validity | - | - | - | - | 01010101 | 01010101 |

Pattern 6F

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 0 | 1 | 1 | 0 | 01100110 | 11111111 |
| Magnitude | 1 | 1 | 1 | 1 | 11111111 | 10011001 |
| Validity | - | - | - | - | 01100110 | 10011001 |

Pattern 7F

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 0 | 1 | 1 | 1 | 01110111 | 11111111 |
| Magnitude | 1 | 1 | 1 | 1 | 11111111 | 11011101 |
| Validity | - | - | - | - | 01110111 | 11011101 |

Pattern AD

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 1 | 0 | 1 | 0 | 10101010 | 01110111 |
| Magnitude | 1 | 1 | 0 | 1 | 11011101 | 10101010 |
| Validity | - | - | - | - | 10001000 | 00100010 |

Pattern DB

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 1 | 1 | 0 | 1 | 11011101 | 11101110 |
| Magnitude | 1 | 0 | 1 | 1 | 10111011 | 01110111 |
| Validity | - | - | - | - | 10011001 | 01100110 |

Pattern AF

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 1 | 0 | 1 | 0 | 10101010 | 11111111 |
| Magnitude | 1 | 1 | 1 | 1 | 11111111 | 10101010 |
| Validity | - | - | - | - | 10101010 | 10101010 |

Pattern BF

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 1 | 0 | 1 | 1 | 10111011 | 11111111 |
| Magnitude | 1 | 1 | 1 | 1 | 11111111 | 11101110 |
| Validity | - | - | - | - | 10111011 | 11101110 |

Pattern DE

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 1 | 1 | 0 | 1 | 11011101 | 10111011 |
| Magnitude | 1 | 1 | 1 | 0 | 11101110 | 01110111 |
| Validity | - | - | - | - | 11001100 | 00110011 |

Pattern DF

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 1 | 1 | 0 | 1 | 11011101 | 11111111 |
| Magnitude | 1 | 1 | 1 | 1 | 11111111 | 01110111 |
| Validity | - | - | - | - | 11011101 | 01110111 |

Pattern EF

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 1 | 1 | 1 | 0 | 11101110 | 11111111 |
| Magnitude | 1 | 1 | 1 | 1 | 11111111 | 10111011 |
| Validity | - | - | - | - | 11101110 | 10111011 |

Pattern FF

|  | Track 0 | Track 1 | Track 2 | Track 3 | Even Channel Pattern | Odd Channel Pattern |
| --- | --- | --- | --- | --- | --- | --- |
| Sign | 1 | 1 | 1 | 1 | 11111111 | 11111111 |
| Magnitude | 1 | 1 | 1 | 1 | 11111111 | 11111111 |
| Validity | - | - | - | - | 11111111 | 11111111 |

# 8	PCMSIM Program

// PCMSIM

// started 970721 pah. This issue  19980128 cleaned up.

//

// Program to simulate PCM operation for a given table and input for PCM

//

// Inputs are :- see usage!

#include <STDIO.H>;

#include <PROCESS.H>;

extern int errno;

int  SMVread( FILE *datfile , int * SMV  );

int  BuildTab(FILE *tabfile , char * ramtab, int LinesToTrash);

int  BuildZeroTab(FILE *tabfile , char * ramtab, int LinesToTrash);

int  CalcSample( long * accums , long sample );

int  LoadTabBytes( char tab[] ,char Databyte );

int  LoadZeroTabBytes( char tab[] ,char Databyte,int BASE );

void usage(void);

char ramtab[32768];     // Ram Table

int SMV[8];             // array of pcal data as addresses to be used.

int main(int argc, char *argv[])

{

int input,output,defaultoutput,InputDataOK,testid,nfields ;

int LinesToTrash,lines;

long accums[9],sample,SamplesToDo;

int SMVgood,i;

FILE *tabfile,*datfile,*zerodat,*results;

char inputline[256];

char *fgetsretval;

if(argc != 6)

{

usage();

exit(1);

}

if ((tabfile = fopen(argv[1], "r"))

== NULL)

{

fprintf(stdout, "Cannot open %s.\n",argv[1]);

exit(1);

}

nfields = sscanf(argv[2],"%d",&LinesToTrash);

if(nfields == 0 ) fprintf(stdout, "fail for nfields\n");

fprintf(stdout, "LinesToTrash %d.\n",LinesToTrash);

nfields = sscanf(argv[5],"%ld",&SamplesToDo);

if(nfields == 0 ) fprintf(stdout, "fail for SamplesToDo nfields\n");

SamplesToDo;

fprintf(stdout, "Table file  %s over %ld samples.\n",argv[1],SamplesToDo);

BuildTab( tabfile , &ramtab[0],LinesToTrash);

fclose(tabfile);

if ((datfile = fopen(argv[3], "r"))

== NULL)

{

fprintf(stdout, "Cannot open %s.\n",argv[3]);

exit(1);

}

else

{

SMVgood = SMVread( datfile , &SMV[0] );

}

fclose(datfile);

if ((zerodata = fopen(argv[4], "r"))

== NULL)

{

fprintf(stdout, "No zero dat file using zero default\n");

}

else

{

fprintf(stdout, "Using zerodata file %s\n",argv[4]);

BuildZeroTab( zerodata , &ramtab[0],LinesToTrash);

}

for(i=0;i<9;i++)

{

accums[i] = 0 ;

}

for(sample = 0; sample < SamplesToDo ; sample++)

{

CalcSample(  &accums[0] , sample );

}

long double ldtemp;

unsigned long ldint ;

for(i=0;i<9;i++)

{

ldtemp = (long double) accums[i];

ldtemp = ldtemp*(long double)((long double)0.25 *(long double)53687.0);

ldint  = ldtemp;

fprintf(stdout, "Accum %d count %ld.  \t0x%08lX\t0x%08lX\n",i,accums[i],accums[i],ldint);

}

return SMVgood ;

}

// Read SMV from data file to construct wave form

// Expected form of input file is

/*

S 01010101

M 00110011

V 00010001

*/

// this will produce a return value of

// VMS

// 000

// 001

// 010

// 111

// 000

// 001

// 010

// 111

int SMVread( FILE *datfile, int *SMVptr  )

{

int SMV[8],S,M,V;

int nfields,bitval,working,i;

char ipline[100];

char Typechar[2],Bits[10];

char *fgetsretval;

S = M = V = 1;

working = 1 ;

for(i=0;i<8;i++)

{

SMV[i] = 0;

}

while( (working) && (S || M || V ) )

{

fgetsretval = fgets(ipline, sizeof(ipline) , datfile);

if ( fgetsretval != NULL )

{

nfields = sscanf(ipline,"%s %s",Typechar,Bits);

if(nfields == 2)

{

switch (Typechar[0])

{

case 'S':

case 's':

bitval = 1;

S=0;

fprintf(stdout, "%s",ipline);

break;

case 'M':

case 'm':

bitval = 2;

M=0;

fprintf(stdout, "%s",ipline);

break;

case 'V':

case 'v':

bitval = 4;

V=0;

fprintf(stdout, "%s",ipline);

break;

default:

bitval = 0;

break;

}

for(i=0;i<8;i++)

{

if(Bits[i] == '1')

{

SMV[i] += bitval;

}

}

}

else

{

working = 0;

}

}

else

{

working = 0;

}

}

for( i=0; i<8; i++) *SMVptr++ = SMV[i];

return 0;

}

long ramtabindex;

int   LoadTabBytes( char tab[] ,char Databyte )

{

int flag;

flag = 0;

tab[ramtabindex++] = Databyte ;

if ((ramtabindex >= 16384 + 3200 ) && (ramtabindex < 20480 )) ramtabindex = 16384 + 4096;

if ((ramtabindex >= 20480 + 3200 ) && (ramtabindex < 24576 )) ramtabindex = 20480 + 4096;

if ((ramtabindex >= 24576 + 3200 ) && (ramtabindex < 28672 )) ramtabindex = 24576 + 4096;

if ((ramtabindex >= 28672 + 3200 ) && (ramtabindex < 32768 )) ramtabindex = 28672 + 4096;

return flag;

}

// Build the Ram table with input from the given table file.

//

int BuildTab(FILE *tabfile , char * ramtab, int LinesToTrash )

{

char text1[100];

int i,InputDataOK,lines, nfields, Dataword,Databyte, Bytes;

char *tabptr, *fgetsretval;

tabptr = ramtab;

//preload all table with zeros

for (i=0;i< 32767 ; i++) *tabptr++ = 0;

*tabptr++ = 0;

ramtabindex = 16384;

InputDataOK = 1 ;

lines = 0;

Bytes = 0;

// Trash all lines down to first real data

while((InputDataOK )&&(lines< LinesToTrash))

{

fgetsretval = fgets(text1, sizeof(text1) , tabfile);

if ( fgetsretval == NULL )

{

InputDataOK = 0 ;

}

else

{

lines++;

}

}

while(InputDataOK )

{

fgetsretval = fgets(text1, sizeof(text1) , tabfile);

if ( fgetsretval == NULL )

{

InputDataOK = 0 ;

}

if ( InputDataOK )

{

if(text1[0] != '#')

{

nfields = sscanf(text1,"%x",&Dataword);

nfields = sscanf(text1,"%x",&Databyte);

if(nfields == 1)

{

LoadTabBytes( ramtab, Databyte );

Bytes++;

}

}

}

}

fprintf(stdout, "Total locations filled %d \n",Bytes);

return 0;

}

int   LoadZeroTabBytes( char tab[] ,char Databyte , int BASE)

{

int flag;

flag = 0;

tab[ramtabindex++] = Databyte ;

if ((ramtabindex >= BASE + 3200 ) && (ramtabindex < BASE + 4096 )) ramtabindex = BASE + 4096;

if ((ramtabindex >= BASE + 4096  + 3200 ) && (ramtabindex < BASE +  8192 )) ramtabindex = BASE +  8192;

if ((ramtabindex >= BASE + 8192  + 3200 ) && (ramtabindex < BASE + 12288 )) ramtabindex = BASE + 12288;

if ((ramtabindex >= BASE + 12288 + 3200 ) && (ramtabindex < BASE + 16384 )) ramtabindex = BASE + 16384;

return flag;

}

// Build the Ram table with input from the given table file.

//

int BuildZeroTab(FILE *tabfile , char * ramtab, int LinesToTrash )

{

char text1[100];

int i,InputDataOK,lines, nfields, Dataword,Databyte, Bytes;

char *fgetsretval;

//preload all table with zeros

ramtabindex = 0;

InputDataOK = 1 ;

lines = 0;

Bytes = 0;

// Trash all lines down to first real data

while((InputDataOK )&&(lines< LinesToTrash))

{

fgetsretval = fgets(text1, sizeof(text1) , tabfile);

if ( fgetsretval == NULL )

{

InputDataOK = 0 ;

}

else

{

lines++;

}

}

while(InputDataOK )

{

fgetsretval = fgets(text1, sizeof(text1) , tabfile);

if ( fgetsretval == NULL )

{

InputDataOK = 0 ;

}

if ( InputDataOK )

{

if(text1[0] != '#')

{

nfields = sscanf(text1,"%x",&Dataword);

nfields = sscanf(text1,"%x",&Databyte);

if(nfields == 1)

{

LoadZeroTabBytes( ramtab, Databyte, 0 );

Bytes++;

}

}

}

}

fprintf(stdout, "Total locations filled %d \n",Bytes);

return 0;

}

// look up in the table rams the data to count

int CalcSample( long * accums , long sample )

{

char sampledata;

long address;

int SMVindex;

// Calculate the address which would be used for the PCal table.

address = sample%3200;

SMVindex = sample%8;

address |= SMV[SMVindex]<<12;

// Get the content for that location.

sampledata = ramtab[address];

// do the accumualtion.

if(sampledata & 0x01) accums[0]++;

if(sampledata & 0x02) accums[1]++;

if(sampledata & 0x04) accums[2]++;

if(sampledata & 0x08) accums[3]++;

if(sampledata & 0x10) accums[4]++;

if(sampledata & 0x20) accums[5]++;

if(sampledata & 0x40) accums[6]++;

if(sampledata & 0x80) accums[7]++;

// Count the validity

if(SMV[SMVindex] & 0x04) accums[8]++;

return SMVindex;

}

void usage(void)

{

fprintf(stdout, "Phase Cal simulator input arguments\n");

fprintf(stdout, "RamTablefile 48 SMVDatafile V=0RamTablefile SamplesCount\n");

fprintf(stdout, "N.B. 48 is the number of lines to trash before first \n");

fprintf(stdout, "Data in the ram table file. \n");

fprintf(stdout, "V=0RamTablefile Defines a file for validity zero if you\n");

fprintf(stdout, "require non zero table data for the validity = zero case.\n");

fprintf(stdout, "There must be an argument. Specify a non existant file\n");

fprintf(stdout, "if you required the default table data = 0 for v=0.\n");

fprintf(stdout, "\n");

fprintf(stdout, "For a single Frame 1:1 there are 6.25 * 3200 samples\n");

fprintf(stdout, "This program assumes mux 4:1 so there are 4 * 6.25 * 3200 samples per Channel frame\n");

fprintf(stdout, "The SamplesCount should be 800000 for above\n");

fprintf(stdout, "The SamplesCount should be 3200000 for 4 frames...\n");

fprintf(stdout, "The Set of numbers on the RHS are for 53687 frames\n");

fprintf(stdout, "assuming the sample count is set for 4 frames.\n");

}